SPECIFICATION FOR DIGITAL MEGA OHMMETER (0-10KV)
· Test Voltages of 500V, 1000V, 2500V ,5000V & 10000V.

· Measures up to min.1 G( in all the ranges with an accuracy of (5% of reading.

· Short circuit current rating max.- 3mA.

· Insulation test voltage accuracy shall be ±5 % .

· Charging rate for capacitive load 5 seconds per (F.

·  Test regime – IR, PI : (In Automatic mode)
· The instrument shall have guard system to eliminate effect of surface leakage current on high resistance measurements.

· It shall have minimum 10 memory locations to store all measurement parameters.

· The instrument shall have warning voltage function to alert the user about presence of line voltage from 30V to 600V ac/ dc.

· It shall have timer for polarization index test (0-30 minutes); settable in 1 minute increments; with least count of 1 second.

· It shall operate on direct Power supply 230 Volts, 50 Hz and also through  high capacity rechargeable batteries through adaptor.

· The instrument shall be supplied with test leads, 10000V rated probes, alligator clips, user’s manual and soft carrying case.

· USB interface for downloading data to PC along with suitable software.

· Megger should be portable & easy to handle.

· Weight with batteries:
Preferably around 5 -7 kgs.

· Temp. range:
Operating:
0° to 40 °C.

Storage:
0° to 50 °C.
· Data Storage : Voltage, Test Time, Leakage Current, Resistance, PI
· Ingress Protection : IP 65

· Other requirements:

a) T/C & G/C including calibration certificate from approved laboratory.
b) O &M manual in English

c) One set of spare leads

d) One spare rechargeable battery

e) Other recommended spares/ toolings

f) PC interface leads

                                            (S.K.DAS)                                                                        (Devanand)




AGM (QC-1)




  Sr.Manager (QC-1)

Other requirement

DIGITAL MEGAOHMMETER (0-10KV)
· Party shall supply two copies of O &M manual in English along with internal circuitry.

· Equipment shall be commissioned at our premises free of cost by the party.

· TC & GC of the equipment are reqd.

· Warranty period shall be of min. one year.

· Calibration certificates are required.

· Party shall supply mandatory spares along with the equipment.

· Party shall offer for pre-inspection of the items before dispatch.

· Party shall submit previous experience of supplying similar equipment or higher rating equipment.

                                (S.K.DAS)

AGM (QC-1)
Pre Qualification requirement

DIGITAL MEGAOHMMETER (0-10KV)
(Indent No.20090786)
· Party shall be regular manufacturer of Digital Mega-Ohmmeter   manufacturer

· Party shall submit previous experience of supplying similar equipment or better specified equipment. Party, whoever does not submit previous P.O. copy (priced /unpriced ) or order reference of at least one party  to whom they have supplied within last 6-7 years, shall not be considered. 

·  Party shall mention the date of commissioning of the supplied equipment to other party.

(S.K.DAS)

AGM (QC-1)

