SPECIFICATION FOR 12 KV,3.0 Amp. AC H.V KIT
· Input Power Requirements: Single phase, 440VAC (±10%), 50Hz.

· Output Voltage

  : 0–12KV AC, 3½ digit Digital Meter,±1% FS Accuracy                                                                          
· Output Current

  : 3000mA, 3 ½digit Digital Meter, ± 2 % FS Accuracy
· Additionally, terminals shall be provided to connect analog meters for measuring output voltage and output current.
· Duty Cycle



: 15 minutes ON/ 15 minutes OFF at full 

  output voltage and current.

· Insulation Class



: F

Equipment shall have following features.

· Manual Controls that is Main Power Circuit Breaker, Main Power On Indicator Lamp, Manual (Push Button type) Raise/Lower of Voltage, Control Power Circuit Breaker or Fuse,  Control Power On Indicator Lamp, Safety Interlock Status Indicator, HV On and Off Pushbuttons, HV On Indicator Lamp, Overload Lamp, Overload Trip Level Switch /  Potentiometer, Overload Reset Push Button.

· In addition to these Auto control features are required that is Digital test Timer, Digital KV meter with Test Voltage Preset and Adjustable Rate of Rise ,Zero inter lock & Auto zero.
· Provision of Caution Light (with flasher) and Audio Buzzer (with volume control) while HV Test system is ‘ON’.

· The equipment shall be supplied with Control Multi-conductor Cables, Instrumentation Coaxial Cables Ground Cables, Input Power Cables, Power cables from the voltage regulator to the test transformer and Voltage Output Cables.
· The complete equipment i.e. HV Test Transformer, Voltage Regulator along with Control Panel having manual controls shall be fixed on a common platform.
· The complete equipment shall be portable. It should be provided with lifting arrangement to shift it from one place to other with the help of crane.
·  Additionally, the platform should be of trailer type for field-testing.

· The readings of the equipment shall be stable and repeatable. 

· Weight of the total equipment shall be less than 1500Kg.

· Ambient operating conditions for the system: 2°C to 45°C and 0-100%RH (Non-condensing).

Party shall indicate other essential spares and optional accessories required for safe operation of the equipment. 

· Party shall supply two copies of O & M Manual in English along with internal circuitry. 

· Equipment shall be commissioned free of cost at our premises. 

· T/C and G/C of the equipment are required.

· Calibration Certificate of HV Test System. 

· Party shall impart training for operation/ maintenance of the equipment. 

· Party shall quote separately for “Annual Maintenance Contract” for 3 years mentioning no. of visits per year.

· Party shall indicate the service facility and service centre in Northern India.

· Party shall indicate self calibrating facility of the equipment, if available (optional).

(S.K.DAS)                                                                                                   (DEVANAND)
Other  requirement

12 KV,3.0 Amp. AC H.V KIT 

· Party shall supply two copies of O &M manual in English along with internal circuitry.

· Equipment shall be commissioned at our premises free of cost by the party.

· TC & GC of the equipment are reqd.

· Warranty period shall be of min. one year.

· Calibration certificates are required.
· Party shall impart training for operation/ maintenance of the equipment while commissioning the equipment to BHEL personnel
· Party shall supply mandatory spares.
· Party shall offer for pre-inspection of the items before dispatch.
· Party shall quote separately for “Annual Maintenance Contract” for 3 years mentioning no. of visits per year.

· Party shall indicate the service facility and service centre in Northern India.

· Party shall submit previous experience of supplying similar equipment or higher rating equipment.

                                (S.K.DAS)

AGM (QC-1)
Pre Qualification requirement

12 KV,3.0 Amp. AC H.V KIT 
( Indent No.20091859)
· Party shall be regular manufacturer of AC HV Kit

· Party shall submit previous experience of supplying similar equipment or higher rating equipment. Party, whoever does not submit previous P.O. copy (priced /unpriced) or order reference of at least one party  to whom they have supplied within last 6-7 years, shall not be considered. 

·  Party shall mention the date of commissioning of the supplied equipment to the other party.
(S.K.DAS)

AGM (QC-1)
