Ref.: EMT/SPEC/1049

|  | BHARAT HEAVY ELECTRICALS LIMITED RAMACHANDRPURAM :: HYDERABAD 502 032. |  | Enquiry No. : |  |
| --- | --- | --- | --- | --- |
|  |  |  | Due Date : |  |
| HYDERABAD | CONTACT DETAILS |  | Supplier Qtn. No.: |  |
|  | DGM/PUR(CG) email : burre@bhelhyd.co.in Phone No. 040-23182250 |  | Date : |  |
| SPECIFICATION CUM COMPLIANCE CERTIFICATION FOR |  |  |  |  |
| DC HIGH VOLTAGE TEST KIT |  |  |  |  |
| NOTE:- |  |  |  |  |
| 1. Vendor (OEM) must submit complete information against clause no. 9. The offer meeting this clause would only be processed. |  |  |  |  |
| 2. The "Offered" Column and where applicable, the "Deviations" & "Remarks" Column of this format shall be filled in by the Vendor and submitted along with the offer. Inadequate / incomplete, ambiguous, or unsustainable information against any of the clauses of the specifications/requirements shall be treated as non-compliance. |  |  |  |  |
| 3. The offer and all documents enclosed with offer should be in English language only. |  |  |  |  |
| ADDRESS OF THE SUPPLIER : |  | ADDRESS OF THE INDIAN AGENTS : |  |  |
| TELEPHONE NOS.: |  | TELEPHONE NOS.: |  |  |
| FAX NOS.: |  | FAX NOS.: |  |  |
| E-MAIL ADDRESS : |  | E-MAIL ADDRESS : |  |  |
| SCOPE: DESIGN, MANUFACTURE, TESTING & CALIBRATION, PACKING AND DISPATCH OF DC HIGH VOLTAGE TEST KIT |  |  |  |  |

| S.No. | BHEL REQUIREMENT | OFFERED | DEVIATIONS | REMARKS |
| --- | --- | --- | --- | --- |
| 1.0 | Purpose of the equipment: |  |  |  |
| 1.1 | The DC high voltage test kit is for use in testing of Generator coils/ half coils and winding during manufacturing. |  |  |  |
| 2.0 | Scope |  |  |  |
| 2.1 | Design, manufacture, testing, calibrate, packing and dispatch of DC High voltage test kit. |  |  |  |
| 3.0 | TECHNICAL REQUIREMENT: |  |  |  |
| 3.1 | Output voltage : 0-50 kV DC (variable) |  |  |  |
| 3.2 | Output current : 5 mA |  |  |  |
| 3.3 | Input voltage : 230 V AC 50 Hz |  |  |  |
| 3.4 | Ripple : less than 2% |  |  |  |
| 3.5 | Regulation : 20% from no-load to full load |  |  |  |
| 3.6 | Duty cycle : 10 minutes ON |  |  |  |
| 3.7 | The test kit shall be provided with ON/OFF controls, circuit breaker for input power supply, supply indication lamp and protective fuses. |  |  |  |
| 3.8 | The test kit shall be provided with ON/OFF controls for High voltage and High voltage ON indication lamp. |  |  |  |
| 3.9 | Continuously adjustable raise/lower control for high voltage shall be provided. The High voltage should not be ON when the raise/lower knob is not in minimum(zero) position. |  |  |  |
| 3.10 | The test kit shall be provided with overload relay. |  |  |  |
| 3.11 | Voltage measurement : Measurement of output voltage by either digital meter or analog meter. The voltmeter shall have resolution for full output voltage measurement or selectable ranges for maintaining resolution and accuracy. The voltmeter shall be protected against damage caused by test sample breakdown. |  |  |  |
| 3.12 | Current measurement : Measurement of output current shall be by either digital meter or analog meter. The ammeter shall have different selectable ranges for maintaining resolution and accuracy. The Ammeter shall be protected against damage caused by test sample breakdown. |  |  |  |
| 3.13 | Accuracy of indicating meters : +/- 2% of full scale. |  |  |  |
| 3.14 | The test kit shall be provided with external permissive interlocking system on the control panel. |  |  |  |
| 3.15 | High voltage components shall be housed in separate sealed module. |  |  |  |
| 3.16 | Control module shall be separate from High voltage module and connected by control cable at least by 15 ft. for safety. |  |  |  |
| 3.17 | The test kit shall be supplied with bleed-off resistor and earthing rod for discharging of charged voltage in the unit. |  |  |  |
| 3.18 | Operating temperature : upto 45 Deg.Cel. |  |  |  |
| 4.0 | Testing : equipment supplied shall be completely tested. |  |  |  |
| 5.0 | Calibration : The complete set shall be calibrated. |  |  |  |
| 6.0 | GAURANTEE : The guarantee of the equipment shall be as per BHEL standard terms and conditions. |  |  |  |
| 7.0 | PACKING and Dispatch : The items supplied shall be packed suitably by vendor so as to exclude any possibilities of damage during transportation and storage at site. |  |  |  |
| 8.0 | Documentation |  |  |  |
| 8.1 | Following shall be submitted with the offer - Technical bid and price bid in separate envelopes - Compliance statement as per the specifications - Technical catalogs. - Reference list of supplies. |  |  |  |
| 8.2 | Following shall be supplied with the supply - Test certificates. - Calibration certificates. - Operation and maintenance manuals. |  |  |  |
| 9.0 | Only those vendors, who have supplied DC High voltage test sets of similar or higher rating (on the Date of opening of tender) and such supplies are presently working satisfactorily for more than one year after commissioning (On the Date of Opening of tender), should quote. However, if vendor had supplied DC High voltage test set to BHEL, then it should be presently working satisfactorily for more than six months after its commissioning and acceptance (On the Date of Opening of tender) in BHEL. Vendor shall furnish reference list of such customers with contact person/s’ details and also order copies. This is required from all Vendors for qualification of their offer |  |  |  |