Annexure (A)

- SCOPE OF SUPPLY:

AUTOMATIC TEST EQUIPMENT (IN CIRCUIT FUNCTION TESTER/ PCB TESTER)  PC based, desk top model working on Window XP platform having following minimum features:

- In-circuit and out circuit testing facility for digital, analogue, mixed signal ICs.

- Analogue devices including OP Amps, Opto coupler, Comparators, Voltage Regulators, PWM controllers etc.

- Mixed Signal Devices including ADCs, DACs etc.

- Programmable Power supply up to 36V/5A.

- 96 Digital Channels.

- 3 analogue channels.

- 8 flying channels for various applications.

- J Tag port for testing boundary scans devices.

- Device library of minimum 25000 no. of devices along with their manufacturer data sheet.

- Test speed of 10 MHz.

- Guarding feature of other ICs during in circuit testing of devices.

- Clock disable feature to eliminate disturbances during in-circuit testing.

- Three level driver state i.e. Hi, Low, Tri- state.

- Built in 2 or more channels Oscilloscope with programmable load.

- VI Signature analysis.

- Identification of Unknown device.

- ROM/PAL/EPROM reading and verification by cell by cell and checksum data.

Essential accessories to be supplied:

- HP/IBM/DELL make desk top PC with minimum configuration Pentium Core 2 Duo 2.66GHz processor, 320GB Hard disk, 2GB  RAM, 16x DVD+/-RW Optical Drive, 9PIN RS232 Port, Preloaded Windows XP, 18.5” Widescreen Flat (TFT) Monitor, USB Keyboard & Mouse.

- Licensed copy of Operating Software and other application softwares with clear written installation instructions.

- User & Maintenance manual in both format i.e. Hard copy as well as on DVD.

- Interface cable for connecting with external controller.

- AC power cord.

- BUT power cable.

- VI Probe & Test Leads as required.

- Micro grabbers to guard channels.

- Standard DIP Test Clips with cables ( 8, 14,16,20,28,40,64 pins 3M make.) -2 no. each.

- SMD SOIC Clip Set consisting of 8P, 14P, 16P, 20P, 24P & 28P Pomona Test Clips for testing SOIC components.-2 no. each.

- PLCC Clips with cable: 20, 28, 32, 44, 52, 68, 84 pin clip with cable. -1 no.each

- Universal Out Circuit Test Socket for SOIC & SOJ packages upto 64 Pins.

- Out Circuit Test Socket for SOP/TSOP/PQFP/PLCC packages upto 64 Pins.

- Out Circuit Test Jig for testing DIP Ics upto 64pins.

- SCOPE OF WORK:

- Installation and commissioning of the AUTOMATIC TEST EQUIPMENT ( IN CIRCUIT FUNCTION TESTER/ PCB TESTER) at CNC Lab. BLK-VI, HEEP BHEL Haridwar.

- Testing of 2 no. of boards having Digital, Analog & Mixed signal ICs in DIP, SOIC, PLCC, SOIC & BGA Packages at our works.

- Testing of  1 no. of board having boundary scan devices using JTAG port at our works.

- Out circuit testing of total 50 no. of Digital, Analog & Mixed signal ICs in DIP, SOIC, PLCC, SOIC & BGA Packages at our works.

- User and Maintenance training to 2 no. of persons at our works for a period of 3days.

3.   TRAINING:

One week training for each to 2 BHEL Engineers in device program development at suppliers works. Traveling, boarding and lodging cost will be borne by BHEL.

4.   WARRANTY:

Vender shall stand warranty for 2 years from final commissioning date against all 		material supplied and work executed.