| SL. No. | PARTICULARS | REQUIREMENT |
| --- | --- | --- |
| 1 | Purpose : | The Digital Automatic Tandelta – Capacitance measurement system is required to conduct Dissipation Factor (Tanδ), Capacitance( C ) test for laboratory’s routine- / investigational- / developmental testing including on Turbogenerator stator winding at site for RLA study as well as to assess the quality of insulation and for establishment of insulation for 800 MW and above rated TGs. |
| 2 | Test Object Details : | The test objects which are to be tested are as follows: - |
| 2.1 | Full Scale Generator Stator Bars : | The rated voltage of generators varies from 10.50 KV to 27.0 KV and capacitance of single generators bars varies from 5000pF to 12000pF. The Dissipation Factor (Tanδ) of a Stator bar is normally measured in steps of 20% of rated voltage and upto 140% of rated voltage. |
| 2.2 | Insulation Samples: | The minimum capacitance of samples is approximately 10 pF and minimum test voltage is approx. 1.0 kV. |
| 2.3 | Stator Winding of Turbo Generator / Hydro Generator : | The rated voltage of generators varies from 10.50 KV to 27.0 KV and per phase capacitance of wound generators varies from 0.2F to 0.6F and maximum capacitance when all phases together is to be tested is approximately 1.8F . The Dissipation Factor (Tan δ) of a Stator winding is normally measured in steps of 20% of rated voltage and upto 140% of rated voltage. So, maximum test voltage will be approximately 40 kV. |
| 3. | Scope of Supply: | The scope of supply includes design, manufacturing, supply, & commissioning and proving of Digital Automatic Tandelta – Capacitance measurement system including basic equipment, standard capacitor and connecting cables etc. at B.H.E.L works and as per technical specification below. |
| 4. | Basic Equipment & Technical Specification : | A modernized state-of-the-art Portable Digital Automatic Tandelta – Capacitance measurement system including suitable Standard Capacitor along with essential accessories to measure the Dissipation Factor (Tanδ), Capacitance of test object as mentioned in clause 2.1 to 2.3 shall have the following major features in it: - |
|  | a ) | For measurement of Tanδ & Capacitance of test objects in ungrounded (floating) as well as grounded test object condition with guard (to nullify the effect on tan delta due to electrically graded portions while measurements are being conducted on a single stator bar as per Clause 8.1.1 to 8.1.4 of IEEE Std 286-2000, & without guard test facility. |
|  | b ) | Standard Capacitor : This shall have the following features : a) Rated Voltage: 50 kV minimum b) Capacitance : 100 pF minimum c) Dissipation Factor (Tandelta) : ≤ 1 * 10-5 d) Partial Discharge level at rated Voltage : ≤ 2 pC |
|  | c) | Integrated state-of-the-art PC with operating software for instrument control, data storage, analysis & report generation. |
|  | d) | Advance software functions like graphical results for trending analysis along with the other important features like digital display of measured value of Capacitance, Tandelta, Leakage Current and Applied Voltage etc.. |
|  | e ) | Built-in standard interface RS 232 / Ethernet. |
|  | f ) | Data storage enabling easy transfer of data to external PC conveniently print or store. |
|  | g ) | Mains cable and other Cable set of length minimum 10 meter suitable for interface of instrument & standard capacitor. |
|  | h) | Operating instructions, calibration & test certificates. |
|  | i) Other Essential Accessories. | External Shunt of 15A or more to be provided in case in-built current capacity of the offered system is less than 15 A. |
|  | j ) Input Supply | Input Supply 230 ± 10% Volt at 50 ± 3% Hz. |
| 4.1 | Measurement Ranges / Accuracy : |  |
| 4.1.1 | For Dissipation Factor (Tan δ) Measurement : | 1. Measuring range of Dissipation factor (Tan δ) : ≈ 0 to 100 (i.e. ≈ 0 to 10000%). 2. Resolution: = 1 * 10-4 or better. 3. Accuracy : (± 0.5% rdg ± 0.0001) or better. Note: In case of higher resolution 1 * 10-5 or better, accuracy ( ± 2% of rdg ± 0.00001) or better shall also be acceptable. |
| 4.1.2 | For Capacitance (C) Measurement : | 1. Measuring range = 1 pF to 5 µF 2. Resolution = 0.01 pF. 3. Accuracy = (± 0.25% rdg ± 0.3 pF) or better. |
| 4.2 | Source of Power Supply & Inter Connecting Cables : | All interconnecting control and power cables shall be supplied by the supplier along with the measuring system. 230V ± 10% Volts at 50Hz ± 3% Hz. power supply shall be made available by B.H.E.L. to the mains of Automatic Capacitance and Tan delta measuring Bridge. Supplier has to ensure operation of the measurement system under above power supply condition. NOTE: - Specifications of all types of required cables and standard capacitor etc. to be included in the Digital Automatic Tan delta – Capacitance measurement system Manual. |
| 4.3 | Spares : | Recommended mandatory spares for minimum 2 years trouble free operation of the Automatic Digital Capacitance and Tan delta measurement system shall be supplied. The details of spares along with its price may please be indicated. |
| 4.4 | O & M Manual : | Operation and maintenance manuals in English shall be supplied with the main Test system. The following details shall be included in operation and maintenance manual |
|  | a ) | Block schematic illustrating the principle of operation of the system. |
|  | b ) | Detailed assembly drawings with clear marking of each assembly /sub- assembly. |
|  | c ) | Installation and commissioning manuals for various components separately and combined. |
|  | d ) | General Maintenance manual of the equipment. |
|  | e ) | Programming manual of controls |
|  | f ) | Specifications for all type of connecting cables for measurement test set-up. |
|  | Note: | Only original copies of O&M manual 2 nos. and not the photo copies to be supplied along with 3 CD ROMs. |
| 4.5 | Calibration Certificates: | Calibration certificate for overall measurement system accuracy & measurement accuracy of all the instruments traceable to National / International standards shall be supplied along with the measurement system. |
| 5. | Packing: | The Digital Automatic Tan delta – Capacitance measurement system / accessories shall be packed in a suitable water proof/vibration proof packing boxes capable of bearing transit hazards. Various packing boxes shall be properly identified and marked giving details of dimensions, net-/ gross weight. The measurement system shall have proper casing/covers for its instruments and assemblies for safe handling and storage. |
| 6. | Commissioning & Proving: | The Automatic digital Tandelta – Capacitance measurement system shall be commissioned by the supplier at B.H.E.L. Any special instrument / equipment required for commissioning shall be brought by the supplier. The Automatic Capacitance and Tan δ measuring system after commissioning shall be proved at B.H.E.L. works. The Digital Automatic Tan delta – Capacitance measurement system shall be considered commissioned only after successful proving. The job / test object required at the time of commissioning shall be provided by B.H.E.L. for proving. |
| 7. | Future Support for Service & Spares : | The supplier shall give an undertaking that even in case of obsolescence of Test System, the party shall give maintenance support to B.H.E.L. for at least 10 years after commissioning with respect to service & spares. |
| 8. | Guarantee : | The equipment / test system shall be guaranteed for a period of at least 12 months after proving at B.H.E.L. works or at least 18 months after shipment. |
| 9. | Optional: | - Prices are to be quoted for each optional item separately, if any. - A large colour built-in TFT touch screen (12” or more). |
| 10. | Training: | Training to 2 ( two ) B.H.E.L. Engineers at BHEL works shall be imparted by the supplier regarding its operation, calibration, maintenance & testing. Supplier shall demonstrate its operation under different test modes ( floating / grounded ) of objects with different test objects having different capacitance value. |
| 11. | General : |  |
| 11.1 |  | Annual maintenance contract, if any, for the measurement system may also be indicated and separately be quoted. |
| 11.2 |  | Any material and / or component not specifically stated above for the complete measurement system but which is necessary for trouble free operation may please be quoted. |
| 11.3 | Compliance Aspect: | The offer shall contain confirmation for compliance to each and every clause of the specification. If there is any deviation, the same shall be clearly stated. |
| 11.4. | Offer Condition: | The offer shall also contain a list of major customers for offered / similar models supplied. |