## Enquiry
BHARAT HEAVY ELECTRICALS LIMITED |  |  |  | * This is only a request for quotation and not an order * The Enquiry No., Date and Due Date shall be super-scribed on the cover containing your offer * We work on Sundays and our weekly-off falls on Thursday until further notice. |  |  |  | ENQUIRY
R35 - 415D | (A Govt. of India Undertaking) Boiler Auxiliaries Plant, Ranipet - 632 406, Tamilnadu, INDIA. | [6 empty] | Enquiry No. and Date |  | Due Date (DD.MM.YYYY)
Fax : +91 - 4172 - 241107 & 241103 Phone: +91 - 4172 - 254343 & 254594 e-mail: mkmoorthy@bhelrpt.co.in (AND) mgs@bhelrpt.co.in | [7 empty] | 2780076E & 15.04.2008 |  | 23.06.2008
Test Certificate |  | Guarantee Certificate |  | Sample | Rev. No. & Date |  | No. of Items | Destination |  | Sheet No.
Required |  | Required |  | Not Required | 0 |  | 1 | BHEL Stores, Ranipet |  | 1
Sl. No. | Material Code, Description & Specification |  |  |  |  | Unit | Quantity | Delivery Schedule
[8 empty] | Lot No. | Delivery Qty. | Dely. Reqd. by
1 | IM8620204700 Air velocity measurement system including velocity sensors, Work station and Data logging software as per the worksheet 'Tech_Spec'. |  |  |  |  | Set | 2 | 1 | 2 | Supplier to specify
Note:
1. This is a Notice Inviting Tender.
2. Sealed tenders (under two part bid) are to be submitted against this Enquiry.
3. Tender Documents - Technical Specification, Commercial Terms & Conditions etc. are to be downloaded from our web site www.bhel.com
4. Any Change/s and / or Corrigenda to the Tender Enquiry including change of tender opening date, if any, would be hosted on the web only. No separate publication in the newspapers or notice through other media would be given to the intending bidders by BHEL.
5. Tender fee for this tender is INR 5000/- (Indian Rupees Five Thousands only) for domestic suppliers and US $ 250/- (US $ Two hundred and fifty only) for foreign suppliers. Tender fee includes all taxes and duties, as applicable.
Supplier's Name and Address |  |  | Dear Sir, | [5 empty] | Please quote Enquiry No., Date, Due Date and Material Code in all your correspondences.
Please submit your lowest quotation on FIRM PRICE, in duplicate subject to our terms & conditions for the above materials so as to reach us on or before the due date by 1000 hrs.
|  |  | 1. Quotations will be opened at 1030 Hrs. on the due date in the presence of the tenderers who may like to be present. | [5 empty] | For and on behalf of Bharat Heavy Electriclas Limited, Ranipet
2. No revision of prices will be entertained after tenders are opened.
3. Late tenders are liable to be rejected.
4. The purchaser shall be under no obligation to accept the lowest or any other tender and shall be entitled to accept or reject any tender in part or full without assigning any reason whatsoever.
|  |  | 5. Manufacture's Name, Trade Mark or patent No., if any should be specified, illustrative leaflets giving particulars are required along with quotation wherever necessary. | [5 empty] | Dy. Manager - Capital Purchase

## Tech_Spec
Boiler Auxiliaries Plant, Bharat Heavy Electricals Limited, Ranipet - 632406, Tamilnadu, INDIA. |  |  | Enquiry No. & Dt.: | 2780076E & 15.04.2008
|  |  | Due Date : | 23.06.2008
Supplier's Ref.:
Date :
Specification cum Compliance Certificate for Air velocity measurement system including velocity sensors, Work station and Data logging software
Note:-
1. The 'Offered' Column and where applicable, the 'Deviations' & 'Remarks' Column of this format shall be filled in by the Vendor and submitted along with the offer. Inadequate / incomplete, ambiguous or unsustainable information against any of the clauses of the specifications / requirements shall be treated as non-compliance.
2. The offer and all documents enclosed with offer should be in English language only.
Name & Address of the supplier: |  | Name & Address of the Indian Agent:
Telephone No. |  | Telephone No.
Fax No. |  | Fax No.
e-mail : |  | e-mail :
Scope: Supply, Installation & Commissioning of Air velocity measurement system including velocity sensors, Work station and Data logging software complying with the following specification
S No. | Description - BHEL's Requirement | Specified | Offered | Deviations | Remarks
1.0 | Scope of supply:
1.1 | Air velocity measurement system shall be capable to measure the low velocity range of ambient air flowing through the physical flow model of Electrostatic precipitator installed in the laboratory. The probe shall house velocity sensors at 5 levels as shown in the enclosed sketch. Measured velocity shall be communicated to the Work station and logged in the system. Suitable application software for data logging shall be developed in LABVIEW and supplied. Custom report shall be generated using the logged data. | Vendor to confirm
2.0 | Probe with velocity sensors specification:
2.1 | Range of velocity to be measured | 0 to 3 m/sec
2.1.1 | Maximum value of velocity measured shall be restricted to 3 m/sec. | Vendor to Confirm
2.1.2 | Least count - details | Vendor to indicate
2.2 | Medium of fluid - Air at ambient temperature | Vendor to note
2.3 | Number of sensors in each probe (sketch enclosed) | 5
2.4 | Probe construction (space between sensors) | 100 mm
2.5 | Length of the probe | 2000 mm (minimum)
2.6 | Diameter of the probe | 25 mm (maximum)
2.7 | Length of the cable (between Probe and system) | 60 metre
2.8 | The sensors shall have high accuracy, fast response to velocity changes and repeatability features in the operating range | Vendor to Confirm
2.9 | Corrosion resistant material shall be utilized in the construction | Vendor to Confirm
2.10 | Yaw and rotational deviations accommodated by the sensors shall be mentioned | Vendor to indicate
2.11 | The sensor electronics supplied shall be lead length independent with RFI, EMI and lightning protection | Vendor to Confirm
2.12 | All wiring shall be shielded and placed in a well-grounded water proof conduit. | Vendor to Confirm
2.13 | Certificate of calibration for the sensors shall be furnished. | Vendor to Confirm
3.0 | Work station Hardware:
3.1 | The offered system shall have the minimum hardware configuration as mentioned below: | Vendor to Confirm
3.2 | Intel Core 2 Duo; 2.66 GHz, DDR RAM 4 GB, 160 GB Hard disc, 21 inch LCD monitor, Optical mouse, Windows key board, LAN Ethernet card. | Vendor to Confirm & indicate details
3.3 | Windows Operating System (XP or latest) + MS Office | Vendor to Confirm
3.4 | A4 size Laser jet printer – 1 no | Vendor to Confirm
3.5 | Ink jet colour printer A3 size – 1 no | Vendor to Confirm
3.6 | Furniture for PC (PC table, PC chair & Printer table) | Vendor to Confirm
4.0 | Data logging software (Application software):
4.1 | The application software shall be built as per the user requirement in LABVIEW. | Vendor to Confirm
4.2 | The application software shall have features to select the sensor one by one, START / STOP buttons for data acquisition. Once the acquisition starts, the data should be displayed in real time in the screen (both analog and digital displays). User shall observe the real time display and log the data at any point of time by pressing ACCEPT button. On acceptance, the data should be logged in real time to the hard disc of the Work station PC. | Vendor to Confirm
4.3 | The application software shall have features of Offline analysis with the logged data like mathematical calculations of the acquired data (details will be provided by BHEL) and comparision of test results with standard requirements. Based on the calculations, graphical representation shall be generated. | Vendor to Confirm
4.4 | Analog and digital display of velocity being measured (output of the selected sensor for logging) shall be made available in the system. Measured velocity shall be displayed in metre / second. | Vendor to Confirm
The system shall also display the following details:
4.4.1 | Minimum, Maximum and Average velocity | Vendor to Confirm
4.4.2 | Standard deviation of the set of logged data | Vendor to Confirm
4.4.3 | All velocity readings shall also be presented in terms of percentage of average velocity. | Vendor to Confirm
4.4.4 | Number of velocity readings above the specified limit shall be presented in both graphical and tabular format. | Vendor to Confirm
4.5 | The user shall have flexibility to perform other analysis like trend plots etc with the software. | Vendor to Confirm
4.6 | A custom report shall be generated from the software with both graphs and tabular data along with other user inputs | Vendor to Confirm
5.0 | Testing, Commissioning & Training:
5.1 | Performance of measurement system shall be witnessed by BHEL before despatch | Vendor to Confirm
5.2 | Vendor shall demonstrate the functionality, install and commission the system at BHEL Ranipet. | Vendor to Confirm
5.3 | Necessary training shall be given to the users for safe and proper use of the system. | Vendor to Confirm
6.0 | Documentation:
6.1 | Five sets of following documents (Hard copies) in English language shall be supplied: | Vendor to Confirm
6.2 | Catalogues, O&M Manuals of all bought out items including drawings, wherever applicable. | Vendor to Confirm
6.3 | Detailed Maintenance manual of the system with all drawings of assemblies/sub-assemblies/parts including Electrical circuit diagrams. | Vendor to Confirm