## Sheet1
BHEL RAMACHANDRAPURAM::HYDERABAD- 502 032
| Description: | YES / NO | Remarks
Single Body Mortuary Cabinet (split model)- Qty 4 Nos.
I | Pre-Qualification of supplier:
1 | Minimum 2 nos. "certificate of satisfactory working" issued by clients
2 | Must have executed min 2 orders of mortuaries each worth Rs.5.0 Lakhs between 31-3-2006 and 31-3-2008
3 | Turnover of the vendor -minimum Rs.10.0 Lakhs for all the last three years.
II | Specification:
1 | Temperature: +2°C to + 5 °C ±1°C, Ambiant +42°C
2 | Capacity- 12000 Btu/hr approx.(or equavalent)
3 | Cooling Cabinet in side size (LxWxH): 8' x 3' x 2' 6" approx.
4 | S.S 304 Stretcher to suit the cooling cabinet with top handles.
5 | Cooling Cabnet should have S.S.angulare sturacture having weighing capacity of 150 kgs. Approx.
6 | Cooling Cabinet shall be grouted to ground & condensing unit to be erected out side the room.
7 | Outer shell- G.I with epoxy & bacteria resistant coating / S.S 304
8 | Interiors- S.S -304 grade
9 | Insulation- 60mm thick Rigid Polyurethane foam (CFC free),Density 40 kg/m³
10 | Top should be transperent un breakable sheet of 5mm thickness
11 | Lid- Flush type(heavy duty) with durable gasket & robust handles.
12 | Lock-with duplicate keys
13 | Vendor shall grout the units to ground in our hospital mortuary room.
III | Refrigeration system:
1 | Heavy duty air cooled Refrigeration system for continuous application
2 | Heavy duty air cooled finned condensor with automatic condensate evaporating system
3 | Non CFC/non-HCFC environment friendly refrigerant gas
4 | Hermatically sealed compressor
5 | Temp Control-Microprocessor based controller
6 | Power supply: 230 V, 1 Ph, 50Hz.
7 | Flexible 3 core power cable (suitable) of 8 mtrs length
Price Bid:
1 | Each Equipment cost X 4 Nos
2 | Installation & Commissioning Charges X 4 nos.
3 | Taxes & Duties X 4 nos.