|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 page 1 of 4 |
| --- | --- | --- |
| 1.0 TECHNICAL SPECIFICATIONS OF STANDARD PRESSURE GAUGES FOR TURBINE OIL GRADE ISO VG-46 - Type : Bourdon - Medium : Turbine oil grade ISO VG-46 - Dial Size : 150 mm - Least Count : 0.1kg/cm² - Working Temperature : 70 Deg C - Accuracy : ±1% of FSD - Mounting : Direct with bottom entry - Pressure Range : as per scope of supply in SECTION-I - Dial Type : Black figure with white background with anti-corrosive finished scale 9. Window : Acrylic/ Laminated safety Glass 10.Enclosure : SS/ die- cast aluminium housing dust proof and splash proof to IP-67 11.Process connection : ¼” NPT 12.Movement : SS-304 Each pressure gauge should be calibrated for Fluid turbine oil grade ISO VG-46. |  |  |

|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 page 2 of 4 |
| --- | --- | --- |
| 2.0 TECHNICAL SPECIFICATIONS OF DIFFERENTIAL PRESSURE GAUGES FOR TURBINE OIL GRADE ISO VG-46 1.Type : Bourdon 2.Medium :Turbine oil grade ISO VG-46 3.Dial Size : 150 mm 4.Least Count : 0.1kg/cm² 5.Working Temperature : 70 Deg C 6.Accuracy : ±1% of FSD 7.Mounting : Direct with bottom entry 8.Pressure Range : as per scope of supply in SECTION-I 9.Dial Type : Black figure with white background with anti-corrosive finished scale 10. Window : Acrylic/ Laminated safety Glass 11.Enclosure : SS/ die- cast aluminium housing dust proof and splash proof to IP-67 12.Process connection : ¼” NPT 13.Movement : SS-304 14.Sensing Element : Diaphragm 15.Diaphragm material : SS-316 16.Diaphragm Chamber : SS-304 17. Max. Stat Pressure : 20kg/cm² Each differential pressure gauge should be calibrated for Fluid turbine oil grade ISO VG-46. |  |  |
|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 page 3 of 4 |
| 3.0 TECHNICAL SPECIFICATIONS OF PRECISION PRESSURE GAUGE INDICATOR - Type : Bourdon - Medium : Air & Hydrogen - Dial Size : 250 mm or more - Least Count : 0.007kg/cm² or better - Working Temperature : 70 Deg C - Accuracy : ±0.07% of FSD or better - Sensitivity : 0.01% of FSD or better - Pressure Range : as per scope of supply in SECTION-I - Dial Type : Black figure with white background with anti-corrosive finished scale 10. Window : Acrylic/ Laminated safety Glass 11. Enclosure : SS/ die- cast aluminium housing dust proof and splash proof to IP-67 Each pressure gauge should be calibrated for Air & Hydrogen. |  |  |

|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 page 4 of 4 |
| --- | --- | --- |
| SECTION-I SCOPE OF SUPPLY The Scope of Supply is as follows: Description; Range; Qty.(nos.). STANDARD PRESSURE GAUGES FOR TURBINE OIL ISO VG-46; 0-6kg/cm²; 20. _________ ,,__________; 0-10kg/cm²; 20. _________ ,,__________; 0-16kg/cm²; 06. _________ ,,__________; 0-250kg/cm²; 06. DIFFERENTIAL PRESSURE GAUGES FOR TURBINE OIL ISO VG-46; 0-2kg/cm²; 04. PRECISION PRESSURE GAUGE INDICATOR; 0-7kg/cm²; 02. |