	[image: image1.png]EAiTELT
‘nﬂnl




	SPECIFICATION OF FLOW METERS REQUIRED AT TEST BED,BLOCK-1
	SPECIFICATION NO. TB-C01-0809
Rev 01

page 1 of  3

	1.0 TECHNICAL SPECIFICATIONS OF ELECTROMAGNETIC FLOW METERS FOR INDUSTRIAL WATER
1. Type                                                : Electromagnetic Flow meter
2. Fluid                                                : Water (Industrial Purpose)
3. Sp. Gravity                                      : 1
4. Pressure Rating                               : 16 bars
5. Flow Velocity                                 : 0.5 m/sec to 10 m/sec
6. Working Temperature                    : 70 Deg C
7. Accuracy                                         : 0.5%
8. Output Transmitter Type                : 2 wire Electrical transmitter inbuilt
9. Transmitter output                          : 4-20 mA DC                                                
10. Main supply                                    : 220 volts AC (±10%), 50 Hz
11. Flow Range                                     : as per scope of supply in 
                                                               SECTION-I
12. Display                                            : Local digital indication of flow rate
13. Mounting                                         : In any of the vertical, horizontal or 
                                                                     Inclined positions
     14.Enclosure protection for

          Electronics                                        : IP 65
          Each flow meter should be calibrated for water.  



	[image: image2.png]EAiTELT
‘nﬂnl




	SPECIFICATION OF FLOW METERS REQUIRED AT TEST BED,BLOCK-1
	SPECIFICATION NO. TB-C01-0809

Rev 01

page 2 of  3

	2.0 TECHNICAL SPECIFICATIONS  OF OIL FLOW METERS FOR TURINE OIL GRADE ISO VG-46
1. Type                                                : Metal tube variable area flowmeter
with Electrical transmitter.
2. Fluid                                                : Turbine oil grade ISO VG-46
3. Sp. Gravity at 50°C                         : 0.85
4. Kinematic viscosity  at 50°C          : 28 CST
5. Pressure Rating                               : 16 bars
6. Working Temperature                    : 70 Deg C
7. Accuracy Class                               : 1.6 or better
8. Output Transmitter Type                : 2 wire Electrical transmitter in-built
9. Transmitter output                          : 4-20 mA DC                                                          
10. Main supply                                    : 220 volts AC (±10%), 50 Hz
11. Flow Range                                     : as per scope of supply in 
SECTION-I
12. Display                                            : Pointer and scale (large scale)
13. Mounting                                         : Vertical         
14. Tube material                                   : Stainless steel

15. Float material                                   : SS 316
16. Flange material                                 : SS 316
17. Enclosure protection for

Electronics                                        : IP 65
18. Flow direction                                   : Bottom to top

19. Flange Size                                        : as per scope of supply

          Each flow meter should be calibrated for Fluid Turbine oil grade ISO VG-46



	[image: image3.png]EAiTELT
‘nﬂnl




	SPECIFICATION OF FLOW METERS REQUIRED AT TEST BED,BLOCK-1
	SPECIFICATION NO. TB-C01-0809

Rev 01

page 3 of  3

	       








SECTION-I

SCOPE OF SUPPLY

The Scope of Supply is as follows:

Description

Flange Size

Range

Qty.(nos.)
ELECTROMAGNETIC FLOW METERS FOR INDUSTRIAL WATER

150 mm

0-150 m³/hr

04
_________ ,,__________

200 mm

0-200 m³/hr

02

_________ ,,__________
200 mm

0-250 m³/hr

02
OIL FLOW METERS FOR TURINE OIL GRADE VG-46

50 mm

0-2 m³/hr

02

_________ ,,__________
50 mm

0-6 m³/hr

04

_________ ,,__________
150 mm

0-100 m³/hr

03




