# Technical Specifications

Digital Ultrasonic Flaw Detector

(Indent No.20084311, Dated 27.10.2008)

| Sl.No. | #### Technical Parameters | # BHEL Specification # Requirement | Relevant technical detail of the offered equipment shall be filled against each row by the bidder. |
| --- | --- | --- | --- |
| 1. | Size of Equipment. | The size of equipment shall not be more than 300mm X 200mm X 180mm. |  |
| 2. | Weight | The weight of equipment shall not be more than 3kg. |  |
| 3. | Display | High resolution LCD / ELD , with Minimum 320 pixels(W)x 240 pixels(H). |  |
| 4. | #### Data Storage | Minimum 200 data measurement and Setup parameters. |  |
| 5. | Transducer connections | 2 X Lemo. |  |
| 6. | Battery operation | Rechargeable Lithium ion, Ni-MH, NiCad battery pack / cells. After full charging the battery should have at least 6 hours working time before recharge, with on screen low battery indication. |  |
| 7. | Power requirement | AC single phase mains, 230 Volts, 50 Hz. |  |
| 8. | Charger | A battery charger compatible with the power supply mentioned above shall be given. |  |
| 9. | Operating Temperature | The equipment shall be capable of working satisfactorily up to the temperature: 5 to 50˚ C. |  |
| 10. | #### Calibration Range | 1 mm to 10000 mm ( Steel ) |  |
| 11. | Sound velocity Range | 1000 metres / second to 15000 metres / second. |  |
| 12. | Frequency Range | 0.5 to 25.0 MHz. |  |
| 13. | Pulse repetition frequency | 30 Hz to 1000 Hz. |  |
| 14. | Gain | 0 to 110 db (Minimum). Minimum two gain steps of 0.10 db and 6db. |  |
| 15. | Pulse Rectification | Full wave, Positive half wave, Negative Half wave, RF |  |
| 16. | Test modes | Pulse echo, Dual, or through transmission. |  |
| 17. | #### Gates | Two fully independent Gates. |  |
| 18. | Reject | Continuously adjustable from 0 to 80% of full screen height. |  |
| 19. | DAC setup | Provision for making Distance Amplitude Correction (DAC) curves on the equipment screen. |  |
| 20. | DGS Evaluation | On screen DGS curve for at least following narrow band, single crystal probes of 2 MHz and 4 MHz frequency, for evaluation of flaws, shall be available: - For Normal beam probes (dia. 24mm) up to 2000 mm range. - For miniature Normal beam probes (dia.10mm) upto 500 mm range. - For Angular beam probes (beam angle 45° , 60° and 70°) of crystal size 20mm X 22mm and crystal size 8mm X 9mm upto 500 mm beam path. |  |
| 21. | I / O ports | The equipment shall have suitable I/O ports for data transfer. |  |
| 22. | Language and units | Menu language shall be English and Units shall be in SI system. |  |
| 23. | Operation/ maintenance manual | At least 3 copies to be provided. |  |
| 24. | #### Guarantee | At least 12 months guarantee against manufacturing defects to be provided from the date of commissioning. |  |
| 25. | Commissioning / training | Equipment is to be commissioned by the supplier at B.H.E.L. works. At least two days training to be provided to UT operators. |  |
| 26. | Operational accessories | A list of spares / probes shall be provided for smooth functioning of the equipment for 5 years. |  |
| 27. | #### Customer List | A list of customers who have been supplied the offered model shall be annexed. |  |