TECHNICAL SPECIFICATIONS OF THREE PHASE TURNS RATIO METER

| Sl. No. | GENERAL DESCRIPTION | Remarks |
| --- | --- | --- |
| 1 | The test set should be fully automatic for finding out the Turns ratio & Voltage ratio of large three phase & single phase Transformers up to 220 KV 315 MVA, Rectifier transformer , CT’s, VT’s |  |
| 2 | The test set should be able to give direct readings of the ratio of test object automatically i.e. microprocessor/PC based without any bridge balancing method involved. |  |
| 3 | The test set should be able to take the measurements automatically. For three phase transformer measurement of the three phases to be conducted simultaneously & automatically without any manual changeover of phases. Phase changeover should be automatic through instrument’s internal circuit. Instrument should have provision for detecting & displaying the following electrical parameters. |  |
| 4 | Display of Turns ratio |  |
| 5 | Display of Voltage ratio |  |
| 6 | Display & detection of vector diagram of transformer under test |  |
| 7 | Display of phase angle deviation |  |
| 8 | The test voltage should be automatic up to the preset value, which shall be determined by the person carrying out the test. |  |
| 9 | The test voltage to be applied on the object under test should be independent of the supply frequency to avoid any variation in the frequency during the testing. |  |
| 10 | The accuracy, reliability, steadiness and the repeatability of the readings should not be affected with the test voltages. |  |
| 11 | Input power supply 230 V +/- 10% , 50 Hz Single phase 2 Amp |  |
| 12 | The test set should have necessary provision for storage of the test results for retrieval and analysis in the associate laptop PC. The required software for Measurements and data storage should also to be supplied along with the test set. The software should be user friendly and Window based. |  |
| 13 | Excitation voltage for Test: 40 volt or higher with user selection criteria Test voltage generation is internal to the instrument |  |
| 14 | Display Back lit graphic display minimum 64X128 Display parameters ratio, excitation current, phase angle deviation |  |
| 15 | Operator control Key pad/ Menu driven for setup, data input, result storage |  |
| 16 | Operating Environment Temperature: 0- 50 C Humidity: 0-90% non condensing |  |
| 17 | Protection: Instrument should be protected against following Transformer high excitation current due to shorted turns/faults Transient voltage protection Surge voltage protection |  |

| The ranges, resolution &accuracy of the test set for different parameters should be at least as per the points mentioned below |  |  |
| --- | --- | --- |
| 1 | Ratio Range – 0.80 to 10,000 Resolution – 4 digit Accuracy – 0.80—1000----- ( ± )0.1 % rdg 1000-2000---- ( ± )0.15 % rdg 2000-5000 ---- ( ± )0.20 % rdg 5000-10,000--- ( ± )0.25 % rdg |  |
| 2 | Excitation Current Range – 0-1 Amp Resolution – 0.1 mA Accuracy – (±) 2.0 % rdg |  |
| 3 | Phase angle Range – + 90 degrees Resolution – 0.1 degree Accuracy – (± ) 0.2 degree or 3 minutes |  |
| 4 | The instrument should have detachable test leads for avoiding damage to the cables as well as to the instrument. |  |
| 5 | The test set should have safety monitoring system to ensure perfect grounding of test system while test is under progress and the accessories required for the same are also to be supplied along with the instrument. |  |
| 6 | The test kit shall meet all the relevant safety specifications |  |

| The Test equipment shall be offered complete with the items mentioned below: |  |  |
| --- | --- | --- |
| 1 | Main equipment--- Transformer Turns Ratio Tester |  |
| 2 | Window based PC software for report generation Report format ratio against tap position in tabular form with all transformer details |  |
| 3 | Power supply cord |  |
| 4 | Cable Set 20 meter single phase set with alligator clip termination 15 meter three phase set with alligator clip termination 10 meter extension set Safety ground lead 5 meter with clamp |  |
| 5 | Two sets of operating and maintenance manuals in English language, |  |
| 6 | Three sets of calibration reports with following data Calibration date, periodicity, calibration validity duration & next calibration due date At least 20 sample calibration results of which minimum 10 results below 100 range Traceability details of standard used for calibration All documents should be endorsed by authorized signature |  |
| 7 | Performance & Warranty Certificate Warranty 2 years after commissioning at our works |  |
| 8 | Storage bag for cables |  |
| 9 | Different type of accessories needed to connect the leads to equipment under testing. |  |
| 10 | Commissioning of the instrument to be done at our work by supplier’s representative free of charge |  |
| 11 | List of critical spare list for 5 yrs trouble free operation to be included |  |
| 12 | One set of spare cable sets to be included in scope of supply |  |
| 13 | Supplier should provide any other attachment/accessories required for smooth functioning of the instrument |  |
| 14 | Original technical catalog to be included along with technical bid |  |
| 15 | Supplier should provide list of critical spares for 5 yrs trouble free operation |  |

TECHNICAL SPECIFICATIONS OF C AND TAN DELTA KIT

| Sl. No. | Description | Remarks |
| --- | --- | --- |
| 1 | The test set should be fully automatic for finding out the capacitance and Tan Delta of large Transformers 220 KV 315 MVA, OIP Bushings, etc in the charged 400kV switchyard. |  |
| 2 | The test set should be able to give direct readings of the capacitance & Tan Delta of the test object automatically i.e. microprocessor/PC based without any bridge balancing method involved. |  |
| 3 | The test set should be able to take the measurements automatically and the results are to be displayed on associated integrated PC/integrated screen and the test set should have provision for displaying the following electrical parameters. |  |
| 4 | Display of Capacitance |  |
| 5 | Display of Dielectric Dissipation factor |  |
| 6 | Display of test voltage |  |
| 7 | Display of frequency |  |
| 8 | Display of Total Insulation Charging Current |  |
| 9 | Display of Inductance |  |
| 10 | Display of Power Loss |  |
| 12 | The test set should be able to suppress the EHV electrostatic and electromagnetic interference by the line frequency modulation method and not by polarity reversal or by injection of opposite interference signal. |  |
| 13 | The test voltage to be applied on the object under test should be independent of the supply frequency to avoid any variation in the frequency during the testing. |  |
| 14 | The test voltage should be automatic up to the preset value, which shall be determined by the person carrying out the test. |  |
| 15 | The accuracy, reliability, steadiness and the repeatability of the readings should not be affected while testing the object for the area involving the 400kV switchyard. |  |
| 17 | The test kit should have provision for providing external interlock for safety to the personnel carrying out the test. |  |

| Sl. No. | Description | Remarks |
| --- | --- | --- |
| 18 | The test kit should be capable of performing tests on all ungrounded and grounded specimens in the following modes without changing the leads anyway. |  |
| 19 | Ungrounded Specimen Test (UST) |  |
| 20 | Grounded Specimen Test (GST) |  |
| 21 | Grounded Specimen Test with Guard (GSTg) |  |
| 22 | The test set should be compatible for the input supply of 230V+/- 10%, 50 Hz+/- 5%, and single phase, AC. |  |
| 23 | The test set should have necessary provision for storage of the test results for retrieval and analysis in the associate laptop PC. The required software for communication. Measurements and data storage should also to be supplied along with the test set. The software should be user friendly and Window based. |  |
| 24 | Power output of the test instrument- Test Voltage: 0-12 kV Power output: 3000 VA (min) At least 100mA continuous at 10kV At least 300mA intermittent at 10kV Frequency—45-70 Hz |  |

| The ranges, resolution, accuracy of the test set for different parameters should be at least as per the points mentioned below |  |  |
| --- | --- | --- |
| Sl.No. | Description | Remarks |
| 1 | The Voltage Range – 80V to 12 kV Resolution – 1 Volt Accuracy – ( ± )1.0 % rdg ± 1 V |  |
| 2 | Capacitance Range – 10 pF – 85.0 nF (At 10 KV & 50 Hz) Resolution – 0.1pF Accuracy – (±) 0.5 % rdg ± 1 pF |  |
| 3 | Tan Delta Range – 0- 100% Resolution – 0.0001 Accuracy – (± )0.5 % rdg ±0.0001 |  |
| 4 | Test Current Range – 30µA -5 Amp Resolution – 0.1μAmp Accuracy – (±) 1.0 % rdg ± 1µA |  |
| 5 | The test set should be able to make the temperature correction in the measurements and the equipment required to measure Temperature and humidity should be supplied along with the test set. |  |
| 6 | The test set should have provision for automatic date & time stamping feature. |  |
| 7 | The test set should be able to operate in an environment of temperature 0 to 50°C & humidity of 10-85% RH (Non condensing) |  |
| 8 | The high voltage unit and the control unit that is the PC should be separate for case of transportation, safe handing and analysis of the results in office by taking out the PC only. |  |
| 9 | The instrument should have detachable test leads for avoiding damage to the cables as well as to the instrument. |  |
| 10 | The test set should have safety monitoring system to ensure perfect grounding of test system and audio/visual alarm while test is under progress and the accessories required for the same are also to be supplied along with the instrument. |  |
| 11 | The test kit shall meet all the relevant safety specifications |  |

| The Specification/configuration of the PC, if necessary should be supplied along with the test set. |  |  |
| --- | --- | --- |
| Sl.No | Description | Remarks |
| 1 | IBM BC compatible (Compaq, Toshiba, Fujitsu or Dell make only) |  |
| 2 | Centrino Duo core Processor |  |
| 3 | 512MB RDRAM, |  |
| 4 | 60 GB hard disk, |  |
| 5 | DVD/CD R/W drive. |  |
| 6 | Battery charger and AC adopter |  |
| 7 | 14” color TFT display |  |
| 8 | AGP2x graphics |  |
| 9 | Built-in modem-56KB and communication lead |  |
| 10 | PCMCIA socket |  |
| 11 | Rechargeable battery for operation of PC |  |
| 12 | Windows latest licensed version |  |
| 13 | Serial port, printer port, USB port |  |
| 14 | Leads- Power-supply, communication, printer (HP make only) etc. |  |
| 15 | Carrying bag for the laptop & Accessories |  |
| The Test equipment shall be offered complete with the items mentioned below: |  |  |
| 1 | Main equipment |  |
| 2 | IBM compatible Laptop PC as per configuration |  |
| 5 | Power supply cord |  |
| 6 | High voltage lead, double shielded of length 20m at least with alligator clip and hook |  |
| 7 | Low voltage leads, shielded of length 20m at least – 3 nos. with necessary clips |  |
| 8 | Ground lead of length 10m with clamp |  |
| 9 | All interconnecting cables & PC interface leads |  |
| 10 | Operating and communication software |  |
| 11 | Three sets of calibration reports with following data Calibration date, periodicity, calibration validity duration & next calibration due date At least 10 sample calibration results of Capacitance & Tan d at different voltage Traceability details of standard used for calibration All documents should be endorsed by authorized signature |  |
| 12 | Performance & Warranty Certificate Warranty 2 years after commissioning at our works |  |
| 13 | Commissioning of the instrument to be done at our work by supplier’s representative free of charge |  |
| 11 | Two sets of operating and maintenance manuals in English language, |  |
| 12 | Tough Carrying case for test leads and accessories |  |
| 13 | Transport case for other attachments, if any |  |
| 14 | Transport Trolley for the main equipment |  |
| 15 | Different type of accessories needed to connect the leads to equipment under testing. |  |
| 16 | One set of spare cable sets to be included in scope of supply |  |
| 17 | Supplier should provide any other attachment/accessories required for smooth functioning of the instrument |  |
| 18 | Supplier should provide list of critcal spares for 5 years trouble free operation |  |
| 19 | Original technical catalog to be included along with technical bid |  |