TECHNICAL SPECIFICATIONS OF RECOVERY VOLTAGE METER

| Sl. No. | Description | Remarks |
| --- | --- | --- |
| 1 | The test set should be fully automatic for finding out the recovery voltage & insulation spectrum of oil paper Insulation system of Transformers, OIP Bushings,CT’s, PT’s. |  |
| 2 | The test set should be able to give direct readings of Insulation resistance of the test object automatically i.e. microprocessor/PC based without any bridge balancing method involved. |  |
| 3 | The test set should be able to take the measurements automatically the test set should have provision for displaying the following electrical parameters. |  |
| 4 | Display of Recovery Voltage |  |
| 5 | Display of Charging Voltage |  |
| 6 | Display of Peak Recovery Voltage |  |
| 7 | Display of time to peak |  |
| 8 | Display of initial voltage rise slope |  |
| 9 | The test voltage should be automatic up to the preset value, which shall be determined by the person carrying out the test. |  |
| 10 | Input Supply : Instrument must be mains/internal battery operated Input supply: Single phase 230 V+/- 10% 50 Hz If battery operated charging circuit should be built in with input supply Single phase 230 V+/- 10% 50 Hz For battery operated instrument permanent on screen display of battery status to be provided |  |

| The ranges of the test set for different parameters should be at least as per the points mentioned below |  |  |
| --- | --- | --- |
| 1 | The Voltage Range – 50-2000 V d.c user selectable Accuracy – ( ± )0.5 % voltage |  |
| 2 | Insulation Resistance Range – 1 M to 1 T Accuracy – (±) 2.0 % rdg |  |
| 3 | Charging & discharging Current Measurement Range:50 mA- 1pA Accuracy: (±) 1.0 % rdg |  |
| 4 | Short Circuit Current--- 200 mA, 100 ms |  |
| 5 | Display: Large Screen LED/LCD display |  |
| 6 | Charging & discharging time range: Tc,Td 10 ms--- 9999 s |  |
| 7 | The test set should be able to operate in an environment of temperature 0 to 40°C & humidity of 10-85% RH (Non condensing) |  |
| 8 | The instrument should have detachable test leads for avoiding damage to the cables as well as to the instrument. |  |
| 9 | The test set should have audio/visual alarm while test is under progress. |  |
| 10 | The test kit shall meet all the relevant international safety specifications |  |

| The Test equipment shall be offered complete with the items mentioned below: |  |  |
| --- | --- | --- |
| 1 | Main equipment |  |
| 2 | Power supply cord |  |
| 3 | Tri-axial measurement cable: 20 m Grounding lead : 10 m |  |
| 4 | Two sets of operating and maintenance manuals in English language, |  |
| 5 | Tough Carrying case for test leads and accessories |  |
| 6 | Three sets of calibration reports with following data Calibration date, periodicity, calibration validity duration & next calibration due date Traceability details of standard used for calibration All documents should be endorsed by authorized signature |  |
| 7 | Performance & Warranty Certificate Warranty 2 years after commissioning at our works |  |
| 8 | Supplier should provide any other attachment/accessories required for smooth functioning of the instrument |  |
| 9 | Original technical catalog with name & address of manufacturer to be included along with technical bid |  |