TECHNICAL SPECIFICATIONS OF POTENTIAL DIVIDER & PEAK KILO-VOLTMETER

| Sl. No. | General Description | Remarks |
| --- | --- | --- |
| 1 | The system i.e. peak kilo-voltmeter with potential divider will be a total integrated system to be used to measure high voltage during power transformer testing. |  |
| 2 | Test to be conducted Power frequency applied voltage test Induced over-potential test with PD measurement |  |
| 3 | System should show direct reading of the voltage with suitable unit without any conversion factor |  |
| 4 | The system should be able to take the measurements automatically the system should have provision for displaying the following electrical parameters. |  |
| 5 | Display of RMS value of voltage |  |
| 6 | Display of peak value of voltage |  |
| 7 | Display of peak/√2 value of voltage |  |
| 8 | System should consist of following modules - HV coupling capacitor for high voltage measurement & input for peak kilo voltmeter - Peak kilo voltmeter for displaying measured value - Necessary inter-connecting cable between coupling capacitor & peak KV meter |  |

| Sl. No. | Specification for Peak KV meter | Remarks |
| --- | --- | --- |
| 1 | The system i.e. peak kilo-voltmeter with potential divider will be a total integrated system to be used to measure high voltage during power transformer testing. |  |
| 2 | Test to be conducted Power frequency applied voltage test Induced over-potential test with PD measurement |  |
| 3 | System should show direct reading of the voltage with suitable unit without any conversion factor |  |
| 4 | The system should be able to take the measurements automatically the system should have provision for displaying the following electrical parameters. |  |
| 5 | Display of RMS value of voltage in absolute value with unit |  |
| 6 | Display of peak value of voltage |  |
| 7 | Display of peak/√2 value of voltage |  |
| 8 | AC mode input voltage from Coupling Capacitor— Voltage Value--0-100 Vrms , Frequency--- 40—200 Hz |  |
| 9 | Over all system accuracy Measured High Voltage range: 0-600 KV Overall system Uncrtainty: +/- 2% of rdg (40-200 Hz) |  |
| 10 | Display: Large Back light LCD display |  |
| 11 | User interface for Setting: Through front panel menu/setting key |  |
| 12 | Following connections to be provided at the rear of the meter 1. AC input 2.Grounding terminal 3.Power supply terminal |  |
| 13 | Meter operating supply Single phase, 230 V + 15%, 50 Hz |  |

| Specification for Coupling Capacitor |  |  |
| --- | --- | --- |
| 1 | Coupling capacitor to be used as direct connection of the measuring system to high voltage point & AC input to Peak kilovolt meter |  |
| 2 | Test to be conducted Power frequency applied voltage test Induced over-potential test with PD measurement |  |
| 3 | Type of use: Indoor |  |
| 4 | Max voltage : 0-600 KVrms, Frequency : 40-200 Hz |  |
| 5 | High voltage end/ top electrode should be of suitable toroidal electrode |  |
| 6 | Base of the coupling capacitor should be fitted with alloy wheel for easy & frequent movement on shop floor in all direction. Wheel & structure should be rugged & suitable for oily surface |  |
| 7 | High voltage capacitor arm should be housed into glass fiber reinforced epoxy tube |  |
| 8 | Suitable termination to be provided for connection between Coupling capacitor & Peak kilo volt meter |  |
| 9 | Suitable termination to be provided for connection between Coupling capacitor & system ground |  |
| 10 | Suitable casing should be provided for handling/shifting of Peak Kilo voltmeter in shop with proper slot for leads connection. |  |

| The Test equipment shall be offered complete with the items mentioned below: |  |  |
| --- | --- | --- |
| 1 | Main equipment— Coupling capacitor & Peak kilo-voltmeter |  |
| 2 | Power supply cord for Peak kilo-voltmeter |  |
| 3 | Interconnecting lead for connection between coupling capacitor & peak kilo-voltmeter with proper termination on both ends—20 mtr |  |
| 4 | Two sets of operation and maintenance manuals in English language. |  |
| 5 | One spare set of Interconnecting lead to be included |  |
| 6 | One set of Peak kilo-voltmeter having identical specifications with necessary leads set to be quoted separately |  |
| 7 | Acceptance routine & type test & performance test as per IEC 60-2 or other relevant standards should be performed on the Measuring system in presence of BHEL representative at supplier’s work. Calibration & test certificates of all the tests performed are to be provided. |  |
| 8 | Performance & Warranty Certificate Warranty 2 years after commissioning at our works |  |
| 9 | Supplier should provide any other attachment/accessories required for smooth functioning of the instrument |  |
| 10 | Original technical catalog to be included along with technical bid |  |
| 11 | Supplier should provide critical spare list for 5 years trouble free operation |  |