- TECHNICAL SPECIFICATIONS OF SINGLE PHASE AC DI-ELECTRIC TEST SET

PURPOSE: this high voltage test system will be used to conduct AC power frequency withstand test of power transformer

A typical test system should include following basic modules

- Regulating Transformer

- Transformer Protection unit

- High voltage transformer

- Potential Divider

- Control  & measuring system

- HV & grounding connection

SPECIFICATIONS OF SYSTEM COMPONENTS

HV TEST TRANSFORMER

- Input Voltage: 0-440V

- Output Voltage: 0- 500 KV

- Variation: Smooth step less

- Output current: 0-2.0 Amp

- Rated Capacity: 1000 KVA

- Duty cycle: 20 min ON, followed by 40 min OFF

- HV transformer to be OIL insulated

- Winding design should take into consideration of high transient voltage withstand capacity during external breakdown.

- Cooling for the transformer should be natural cooling (ONAN)

REGULATING TRANSFORMER

- Capacity of the regulator should be adequate to support  power requirement of total test system

- Regulator Input Voltage: 440 Volt

- Regulator Maximum Input current:

- Output Voltage (input to HV transformer): 0—440 Volt

- Voltage variation: Smooth step less

- Voltage variation method: Using motor control through RAISE & LOWER push button integrated to control module

- Regulating transformer should be oil/air cooled with natural mode of cooling

- Regulator should have input breaker & output breaker/isolator

- Input breaker & output breaker/isolator operations to be controlled remotely from control desk

- Primary breaker & secondary breaker/isolator should be placed in regulator cabinet

TRANSFORMER PROTECTION UNIT & SAFETY FEATURES

- Controls in a metal enclosure, with provision for a separate ground lead.

- External interlock provision

- Surge/transient protection on meters and transformers

- Overload circuit adjustable from 10% to 110% of rated current; includes indication with reset.

- Main Power Circuit breaker

- Latching emergency off mushroom pushbutton

- Zero-start interlock

- Backup overload circuit in the primary of the high-voltage transformer.

POTENTIAL DIVIDER

- Potential Divider  should be connected to high voltage terminal for voltage measurement purpose

- Output of capacitor to be provided to KV meter on control desk.

CONTROL & MEASURING SYSTEM

Total control & MEASURING system should be clustered on a single desk type cabinet

Control system should consists following

Switching On/Off of the primary circuit breaker

Switching On/Off of the contactor (High Voltage On/Off)

Manual variation of the output high voltage

Automatic zero start-up after switching off the high voltage

Timer for switching off the high voltage after a preset time

At any time and in addition to the output high voltage a selection of four of the following information can be displayed:

Status of the protection system

Status of the safety circuit and the emergency trip

Output voltage and current of the regulating transformer

Timer

Measuring system should consists following

- The measuring signal is derived from the HV divider. The value of the high voltage is permanently displayed, either as peak value/2 or as RMS value.

- The voltage display is in digital form with 31/2 digit bright LED/ backlit graphic LCD display

- Analog/digital type ammeter for measurement of leakage current

- Flashover at the test object the control unit automatically stores the last voltage measurement.

- Measuring error of the complete voltage measuring1 % of output voltage at 50 / 60 Hz.

- Measuring error of the complete current measuring system is  2 % of output voltage at 50 / 60 Hz.

HV & GROUNDING CONNECTION

- Ground studs to be provided on the controls, voltage regulator, coupling capacitor and high voltage transformer for permanent connection to a station ground.

- Supplier should provide all necessary cables for connection between different modules

STANDARD AMBIENT CONDITION OF AC TEST SYSTEM

- Height above sea level <1000 m

- Relative humidity in main hall under non condensing conditions < 90 % non condensing

- Temperature averaged over 24 h for H.V. components min. 5 °C, max. + 40 °C

- Extreme temperatures for H.V. components min. 0 °C max. + 50 °C

DOCUMENTATION

Three sets of following document s (Hard copies) in English language should be supplied along with system

- Operating manual of the system

- Maintenance manual with all drawings & electrical circuit diagram

SYSTEM INSPECTION

- Inspection at supplier works before dispatch

- Tests/activities should be carried out at supplier’s work as per relevant international standards before dispatch in presence BHEL representatives

- Demonstration of all controls, operation & measurement performance

SYSTEM INSTALLATION & COMMISSIONING

- Supplier should take full responsibility for carrying out the start up & commissioning

- Commissioning activity should include operational test to full capacity of the system &  on site training on BHEL staff on functioning, control of the system

- GENERAL CONDITIONS

- OGA drg approval is required from bhel prior to manufacturing

- Electrical scheme approval is required from bhel prior to manufacturing

- Only vendors having experience of supplying installing and commissioning of eqipts similar or higher rating will qualify to quote for this tender. Relevent documents shall be submitted alongwith the offer.