TECHNICAL SPECIFICATIONS OF DUAL CHANNEL MICRO OHMMETER

(COMPLIANCE SHEET)

| Sl.No | Description | Remarks |
| --- | --- | --- |
| 1 | The test set should be fully automatic for finding out the winding resistance of three phase & single phase Transformers up to 220 KV 300 MVA rating & Rectifier transformer of 100 KA current rating. Instrument should be suitable to provide stable resistance value of high inductive winding quickly. |  |
| 2 | The test set should be able to give direct readings of the resistance of test object automatically i.e. microprocessor/PC based without any bridge balancing method involved. |  |
| 3 | No of channel—2 (Dual Channel) |  |
| 3 | Measurement Principle—Four Wire Kelvin Bridge |  |
| 4 | Simultaneous Display of Resistance Value of both channel |  |
| 9 | The test current should be automatic up to the preset value, which shall be determined by the resistance value or measurement range selected under test |  |
| 10 | The test voltage to be applied on the object under test should be independent of the supply frequency to avoid any variation in the frequency during the testing. |  |
| 11 | The accuracy, reliability, steadiness and the repeatability of the readings should not be affected with the test voltages. |  |
| 12 | Input power supply 230 V +/- 10% , 50 Hz Single phase |  |
| 13 | Test Current Auto range 25 Amp or higher & continuous |  |
| 14 | Excitation voltage for Test 50 volt dc or higher Test voltage generation is internal to the instrument |  |
| 15 | Display Back lit graphic display with large font or red LED Display resistance value with appropriate unit |  |
| 16 | Operator control Key pad/ Menu driven for setup, data input, result storage |  |
| Protection: Instrument should be protected against following |  |  |
| Sl.No | Description | Remarks |
| 1 | Discharge the specimen safely when test is completed, if lead accidentally disconnects or if power is lost |  |
| 2 | Monitoring circuit for contact operation of on load tap changes for proper make before break sequence. If open circuit exist instrument should shut down immediately without any hazard/damage to instrument or operating personal. |  |
| 3 | Both visual & audio Hazard warning should be provided for following Test in progress Test specimen is discharging Test specimen is fully discharged ready for connection changing |  |

| The ranges, resolution, accuracy of the test set for different parameters should be at least as per the points mentioned below |  |  |
| --- | --- | --- |
| 1 | Resistance Range – 1 m to 100  Resolution – 4 digit Accuracy – ( ± )0.1 % rdg |  |
| 2 | Current Range – 25 Amp or higher & continuous Auto range |  |
| 3 | Stability : Reading should be stable & in steady state variation in measured value shouldn’t be more than +/-0.1% of mean value |  |
| 4 | The instrument should have detachable test leads for avoiding damage to the cables as well as to the instrument. |  |
| 5 | The test set should have safety monitoring system to ensure perfect grounding of test system while test is under progress and the accessories required for the same are also to be supplied along with the instrument. |  |
| 6 | The test kit shall meet all the relevant safety specifications |  |

| The Test equipment shall be offered complete with the items mentioned below: |  |  |
| --- | --- | --- |
| 1 | Main equipment |  |
| 2 | Window based PC software for remote operation, data collection, report generation & transformer heat run testing for curve plotting to determine resistance value at t=0 Report format resistance against tap position in tabular form with all transformer details & temperature conversion from measuring temperature to standard temperature or embedded software for heat run calculation. Heat run graph plot & winding rise calculation report. |  |
| 3 | Power supply cord |  |
| 4 | Cable Set 20 meter potential lead with proper termination for resistance measurement—2 pair 20 meter current lead with proper termination (C clamp termination)--- 1 pair 5 meter winding inter connection lead—1 no Safety ground lead 5 meter with clamp |  |
| 5 | Two sets of operating and maintenance manuals in English language, |  |
| 6 | Three sets of calibration reports with following data Calibration date, periodicity, calibration validity duration & next calibration due date 20 sample calibration results of which minimum 2 results at each resistance range Traceability details of standard used for calibration All documents should be endorsed by authorized signature |  |
| 7 | Performance & Warranty Certificate Warranty 2 years after commissioning at our works |  |
| 8 | Storage bag for cables |  |
| 9 | Different type of accessories needed to connect the leads to equipment under testing. |  |
| 10 | Commissioning of the instrument to be done at our work by supplier’s representative free of charge |  |
| 11 | One set of spare cable sets to be included in scope of supply |  |
| 12 | Supplier should provide any other attachment/accessories required for smooth functioning of the instrument |  |
| 13 | Supplier should provide list of critical spares for 5 years trouble free operation |  |
| 14 | Original technical catalog to be included along with technical bid |  |
| Note: Supplier should submit his technical compliance/comments against each of the above points in above format. Technical bid without complete technical compliance sheet will not be considered for technical evaluation |  |  |