TECHNICAL SPECIFICATION NO : QC-IV-79511 ( REV. 00)

TECHNICAL SPECIFICATION OF HELIUM LEAK DETECTOR.

1. PURPOSE:-

Helium Leak Detector is required to check the brazing leakage in water boxes used in 500/800MW TG Stator bars.

2. SCOPE OF SUPPLY:-

The scope if supply include design, manufacturing, supply & commissioning and proving of Helium Leak Detector at B.H.E.L. works and as per technical specification below.

3. TECHNICAL SPECIFICATION:-

  - TEST OBJECT DETAILS:

500/800 ME TG stator bars of rated voltage from 21.0  to 27.0 KV are to be checked 100% for helium leak test after brazing of water boxes and thermal shock test.

4. BASIC EQUIPMENT & SPECIFICATION:-

Helium Leak Detector is required to function in SNIFFING mode. It shall be portable and dry (oil free roughing system) with following parameters:

* Measuring Range:		1x10-7 to mbar. 1/s

* Sensitivity:			1x10-7 to mbar. 1/s

* Detectable Gas:		Helium 4

* Response time:		<1 second with 5 m sniffer probe length.

* Clean up time: 		< 10 second after sniffing Helium.

* Start up time:		< 5 minutes

* Roughing system:		Diaphragm pump + Vane pump / Turbo pump + Scroll

Pump (Oil free)

* Power supply:		230 V, 50Hz

Integrated Functions:

- Auto calibration

- Auto zero

- Automatic signal correction

- Auto alarm

- Hand held remote control

- Interface board RS 232

Miscellaneous Requirements:

- Test Certificates are required with the supply.

- Mandatory spares shall be quoted with the offer.

4.1 SPARES:-

Recommended mandatory spares for 2 years trouble free operation of the Helium Leak Detector shall be supplied with main equipment. The details of spares along with its price may please be indicated.

4.2 O & M MANUAL:-

Operation and maintenance manuals in English shall be supplied with the equipment. Only original copies of O & M manual 2 nos. and not the photo copies to be supplied along with 3 CD ROMs.

4.3 CALIBERATION CERTIFICATES:-

Calibration certificate of Helium Leak Detector to be sent along with equipment.

5.0 PACKING:-

The Helium Leak Detector shall be packed in a suitable water proof/vibration proof packing boxes capable of bearing transit hazards. Various packing boxes shall be properly identified and marked giving details of dimensions, net/gross weight. The equipment shall have proper casing/ covers for its instruments/assemblies for safe handling storage.

6.0 COMMISSIONING & PROVING:-

The Helium Leak Detector shall be commissioned by the supplier free of cost at B.H.E.L. works. Any special instrument/equipment required for commissioning shall be brought by the supplier.

The equipment after commissioning shall be proved at B.H.E.L. works. The equipment shall be considered commissioned only after successful proving. The job/test object required at the time of commissioning shall be provided by B.H.E.L. for proving.

7.0 FUTURE SUPPORT FOR SERVICE & SPARES:-

The supplier shall give an undertaking that even in case of obsolescence of Helium Leak Detector, the party shall give maintenance support to B.H.E.L. for 7 years after commissioning with respect to service & spares.

8.0 GURANTEE:-

The Helium Leak Detector shall be guaranteed for a period of 12 months after commissioning at B.H.E.L. works.

9.0 OPTIONAL:-

Prices are to be quoted for each optional item separately, if any.

10.0 TRAINING:-

Training of 2 B.H.E.L. engineers at B.H.E.L. works in the area of quality control, maintenance & testing shall be imparted by the supplier. Supplier shall demonstrate its operation.

11.0 GENERAL:-

11.1 Annual maintenance contract for the Helium Leak Detector.

11.2 Any material and/or component not specifically stated above but which are

necessary for trouble free operation of the equipment and its accessories.