SPECIFICATION FOR IMPULSE TESTER

TEST VOLTAGE:


0 to 12 KV (Continuously Variable)

PULSE ENERGY:


100 Joules (Approx.). at 12 KV.

OUTPUT (ON NO LOAD):

Nominal


Maximum






Waveform (( sec.)

Voltage (KV)

0.2 / 5



12
DISPLAY & MEASUREMENT

OF OUT PUT VOLTAGE        :
Digital storage oscilloscope with software for 

                                                       Printing of graphical representations. 

 31/2 digit Digital voltmeter to indicate the    

  charging voltage.

CONTROLS:



MCB - for Mains ON/OFF.

Push Button for HT ON./OFF
Manual /Auto trigger switch.

Test voltage control knob.

LIGHT INDICATOR:

Control unit ‘ON /OFF‘ Indication lamp.






Green lamp indication for output ground.

DIGITAL STORAGE OSCILLOSCOPE

Bandwidth:



100 MHZ

Sample rate:



500 MS/S 

Time/Division:



5 ns to 50 s/div.

Channel:



Two Channel 

Trigger:



External, Internal

Extra backlit display.

Advance Trigger:



Delay, Pulse  & Video.

Supply:



Rechargeable Battery cum 230 V Mains supply 

Computer interface:



USB and RS-232 Port for connecting to PC.

The following accessories are to be provided along with the equipment :

1. Connecting Leads–
a) Impulse tester to Oscilloscope.

b) Oscilloscope  to  printer.

2.  Re-chargeableBattery, Fuses & any other recommended spare.

· Other requirements:

a) T/C  & G/C 

b) O &M manual in English

c) One set of spare leads

d) Other recommended spares/ toolings

e) PC interface leads

· The equipment shall be mounted on a dust free rack mounted trolley with four castor wheels for easy towing on shop floor.   

                                    (S.K.DAS)                                                        (Devanand)




AGM (QC-1)




  Sr.Manager (QC-1)

Other  Requirement

IMPULSE TESTER
· Party shall supply two copies of O &M manual in English along with internal circuitry.

· Equipment shall be commissioned at our premises free of cost by the party.

· TC & GC of the equipment are reqd.

· Warranty period shall be of min. one year.

· Calibration certificates are required.

· Party shall supply mandatory spares along with the equipment.

· Party shall offer for pre-inspection of the items before dispatch.

· Party shall submit previous experience of supplying similar equipment or higher rating equipment.

                                (S.K.DAS)

AGM (QC-1)

Pre Qualification requirement

    IMPULSE TESTER
(Indent No.20090777)
·   Party shall be regular manufacturer of Impulse Tester   manufacturer

· Party shall submit previous experience of supplying similar equipment or better specified equipment. Party, whoever does not submit previous P.O. copy (priced /unpriced ) or order reference of at least one party  to whom they have supplied within last 6-7 years, shall not be considered. 

·  Party shall mention the date of commissioning of the supplied equipment to other party.

(S.K.DAS)

AGM (QC-1)

