## specs
BHARAT HEAVY ELECTRICALS LIMITED
| Heavy Power Equipment Plant, Ramachandrapuram, |  | M209CG8P0276BDP
Hyderabad - 502 032. INDIA.
Specification cum Compliance certificate for THE SUPPLIER FOR SUPPLY OF Power & Energy meter calibrator , 1 No.
Note: 1. Vendor must submit complete information against each clause. 2. The Vendor should fill " Offered" column in compliance to specified requirement and also "Deviations" Column, where there is a deviation from the requirement. Duly filled in specification cum Compliance certificate shall be submitted along with the offer. Inadequate, Incomplete , Ambiguous, or unsustainable information against any one of the clauses of the specification/ requirement shall be treated as non-compliance. 3.The offer and all documents enclosed shall be in English Language only.
| ADDRESS OF THE SUPPLIER : | ADDRESS OF THE INDIAN AGENTS:
| TELEPHONE NOS.: | TELEPHONE NOS.:
| FAX NOS.: | FAX NOS.:
| E-MAIL ADDRESS : | E-MAIL ADDRESS:
Sl. No. | Description for BHEL Requirement | Offered | Deviations | Remarks
1 | Purpose :The equipment shall be of table top & usable to calibrate Analogue, Digital Power meters, Power Transducers and Energy meters.
2 | Technology :Equipment shall be of state of art technique with Digital signal processor based system.
3 | Capability to automate the process of calibration using PC and software
4 | Operation : Stand alone equipment shall be operable from the front panel keys and the shall contain the display of all the parameters set/ selected.
5 | Voltage & Current setting resolution : 6 digit
6 | Single phase AC
7 | Voltage range :1 to 700 V, continuously variable in the range
8 | Accuracy : 100 to 150 ppm
9 | Burden / Compliance on Voltage output : > 50 mA
10 | Current range: 0.05 A to 80A, continuously variable in the range
11 | Current accuracy :140 to 700 ppm
12 | Burden / compliance on Current output : > 10 V
13 | Frequency: Range 20Hz to 400Hz
14 | Frequency accuracy : 50 ppm
15 | frequency setting resolution : 0.1 Hz
16 | Phase setting range : -180 to +180 °
17 | Resolution : 0.001 °
18 | Accuracy on Phase setting: 0.003 to 0.015 °
19 | Power setting Range : 3 to 37,500 VA continuously variable in the total range
20 | Power accuracy : 250 to 500 ppm @upf 300 to 500 ppm @ 0.75 > pf >0.5 300 to 600 ppm @ 0.5 > pf > 0.25
21 | Shall have capability to calibrate Single Phase Energy meter with current & Voltages ranges. Shall have capability to give ENERGY set value to the Energy meter under calibration. Shall have capability of calculating energy taken to make a one revolution by Meter under calibration.
22 | The Energy output accuracy : Better than 500 ppm of set value.
23 | The equipment shall have capability of introducing Flicker, Harmonics, DC, dip,swell
24 | Power supply requirement : 230 ± 10 %, AC, 50 Hz, Single Phase
25 | Operating temperature : 5 to 35 °C
26 | The effect of Ambient temperature on performance / output parameters shall be given in the offer
27 | The equipment shall be robust to with stand the transportation (shock and vibration) for the purpose of Calibration at other Calibrating agency
28 | Calibration Certificate for the equipment supplied shall be from NABL accredited Calibration lab/ or from any Accredited lab under ILAC or APLAC bodies for its full specification for all parameters (a) Voltage, (b) Current , (c) Power, (d) Power Factor. (Shall be done after placement of Order by BHEL)
29 | Soft ware to use the equipment using PC for programme assisted calibration of Power meters, Energy meters, voltmeter, ammeter. Full details of software shall be enclosed with offer. The software offered shall be of latest upto date released by the company. Any update released with in warrantee period (the update) shall be given at free of cost.
30 | Accessories -essential for its operation shall be quoted item wise with full details of utility & functions
31 | Accessories optional shall be quoted with details item wise including functions and utility
32 | Weight of equipment shall be quoted in the offer
33 | Documents & manuals:Shall supply (a) User operational manual, (b) Calibration manual, (c) 5 sets of Service & Maintainance manuals containing all assembly /sub assembly drawings, circuit drawings of machine etc in english language.d) The manual also should contain detailed spare parts list along with specification,part number and address of manufacturer.
34 | Spares for maintenance of equipment if any, with break up shall be quoted as separate item
35 | Commissioning :The supplier shall commission the equipment at BHEL, Hyderabad at free of cost and demonstrate the all features and operation of Equipment
36 | Acceptance of Equipment: The equipment will be accepted by BHEL upon satisfactory demonstration of all its features and performance
37 | The training on usage of equipment, software supplied shall be given at on sight by the Indian Agent to BHEL personnel at free of cost.
38 | Qualifying condition: Only the vendors Supplied and Commissioned the same model with in India and running satisfactory for at least one year. (Order copy, credentials of Customer and full details of customer for contact shall be provided with the offer)
39 | Annual Maintenance Contract (AMC) charges per year for 3 years after the warrantee period shall be quoted as separate item
40 | Warrantee : The supplier shall provide warranty for the equipment as per the commercial terms and conditions of BHEL purchase.
41 | Packing: Rigid packing to be made for all items of complete machine, all accessories and other supplied items to avoid any damages/loss in transit. Whenmachine is despatched.All small loose items shall be suitably packed in boxes.Packing list to be provided in the boxes