SPECIFICATION FOR SOUND ANALYZER

1.0  General Description :

This specification covers  the requirement for a compact ,portable digital sound level analyzer capable of

measuring and analyzing real time frequency in 1/1 and 1/3 octave bands.

- Technical  Specification :

2.1	      Measurement range        : 25 - 140 dB

2.2      Frequency Range             :   20 – 20000   Hz

2. 3     Frequency Weighting      :  A , C ,  Z

2.4      Time weighting                : Fast ,Slow and Impulse ; all  conforming to IEC 61672 –

Class I. Peak detector < 50/100 μS.

2.5      Analysis                            :  Real time frequency analysis in 1/1 and 1/3 octave band

filter selection through PC based software modules.

2.6     Display                              :   Back lit color LCD Display

2.7     Measurements                   : Display only : Lpeak , Lp for A,C or Z weighting.

:  Display & Store : Lpmax, Lpmin,Leq, real time and Date.

2.8     Memory                            : Minimum 2 MB internal memory

2.9     Microphone                      : 1) Type-1 Pre polarized ½” condenser

2) Type-1 Pre polarized 1/4” condenser

2.10   Microphone Pre-amplifier : Compatible amplifiers with above microphones.

2.11  Measurement Annotation  : Voice /text

2.12  Software Support               :  Facility to store, down load and analyze real time audio frequency using

compatible latest windows based  software (to be supplied along with the

Equipment) through  USB  port.

2.13   Environmental                  : Operating temp : 5 – 50  Deg C; Humidity : upto 95% RH

2.14  Applicable Standard         : IEC 60651 : 1979 Type-I

IEC 60804 : 2000 Type-I

IEC  61672  : 2002 Class I

2.15   Power supply                   : Battery as well as mains ( 230 V AC , 50Hz) ;

Rechargeable battery with charger : Battery should be

Suitable for 8 hours continuous operation.

Note :  1)  Warranty period : 18 months from the date of commissioning or 24 months from

the date of supply.

- The sound level analyzer should be commissioned and demonstrated for proper

Operation and functioning at BHEL works.

3 ) 3 sets of O&M manuals are to be supplied .

4)  Training for operation of sound level analyzer shall be provided by the vendor at

BHEL works for  2 days.

RECORD OF REVISIONS

| REV.NO | DATE | REVISION DETAILS | REVISED | APPROVED |
| --- | --- | --- | --- | --- |