| BHEL |  |  | # TECHNICAL SPECIFICATION FOR digtal hand held instrumnets |  |  |  | Ref. : GT/SPEC/05/1014 Rev. : 00 Page 1 of 3 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Intent of specification : This specification intended to cover the design, manufacturing, assembly, testing at manufacturer works, delivery of the instrument/s at site in properly packed and forwarded with all accessories as specified in the scope of work and as required for the safe and trouble free operation of equipment at site. Application of System : These instruments will be used to measure the ambient condition (pressure, temperature, humidity and airflow etc.) Scope of Supply : One instrument each. Technical Detail : 1. Environmental Condition : - Ambient Temperature : 15- 50ºC - Ambient Pressure : 710-730 mm of Hg - Relative Humidity : 95% Max. - Out door installation. 1. Hand Held Digital Pressure Indicator : |  |  |  |  |  |  |  |  |
|  | Sl.NO. | Parameters |  |  | Details |  |  |  |
|  | 1. | ### Range |  |  | 0 to 1.25 Bar Absolute |  |  |  |
|  | 2. | ### Accuracy |  |  | ± 0.1% Full Scale including non-linearity, hysteresis and repeatability |  |  |  |
|  | 3. | ### Units |  |  | Bar, mBar, Kgf/cm² , mm of H2O , mm of Hg, PSI , Kgf/m² etc. |  |  |  |
|  | 4. | ### Pressure port |  |  | 6 mm OD hose connector |  |  |  |
|  | 5. | ### Max/ Min/Hold |  |  | Capture minimum, maximum values and holding the values |  |  |  |
|  | 6. | ### Zero |  |  | Zero Correction |  |  |  |
|  | 7. | ### Power Supply |  |  | Battery with low battery indication and 220V, 50Hz AC with adopter |  |  |  |
|  | 8. | ### Display |  |  | Back Lit LCD Display |  |  |  |
|  | 9. | ### Scope of Supply |  |  | Instruments with user manuals and calibration certificates. |  |  |  |
| Date : 15-04-2005 |  |  |  | Prepared by : |  | Approved by : |  |  |

| BHEL |  |  | # TECHNICAL SPECIFICATION FOR digtal hand held instrumnets |  | Ref. : GT/SPEC/05/1014 Rev. : 00 Page 2 of 3 |  |
| --- | --- | --- | --- | --- | --- | --- |
| 2. Hand Held Digital Hygrometer and Temperature indicator: |  |  |  |  |  |  |
|  | Sl.NO. | Parameters |  | Details |  |  |
|  | 1. | ### Range |  | 10% - 95% RH -20ºC to 60ºC |  |  |
|  | 2. | ### Resolution |  | ± 0.1% RH , 0.1ºC / 0.1ºF |  |  |
|  | 3. | ### Accuracy |  | Humidity ±3% RH , Temperature ±1ºC |  |  |
|  | 4. | ### Units |  | Humidity - % RH Temperature - ºC , ºF |  |  |
|  | 5. | ### Max/ Min/Hold |  | Capture minimum, maximum values and holding the values |  |  |
|  | 6. | ### Power Supply |  | Battery with low battery indication and 220V, 50Hz AC with adopter |  |  |
|  | 7. | ### Display |  | Back Lit LCD Display |  |  |
|  | 8. | ### Scope of Supply |  | Instruments with user manuals and calibration certificates. |  |  |
| 3. Hand Held Digital Air Flow meter : |  |  |  |  |  |  |
|  | Sl.NO. | Parameters |  | Details |  |  |
|  | 1. | ### Range |  | Air Flow : 0 – 2500 m³/hr Velocity : 0 – 40 m/sec |  |  |
|  | 2. | ### Resolution |  | Flow : 1 m³/hr Velocity : 1m/sec |  |  |
|  | 3. | ### Accuracy |  | 3% |  |  |
|  | 4. | ### Type |  | Rotating Vane type , Sensor Dia 22 mm |  |  |
|  | 5. | ### Units |  | Flow : m³/hr , l/sec , m³/sec , CFM Velocity : m/sec, ft/sec |  |  |
|  | 6. | ### Max/ Min/Hold |  | Capture minimum, maximum values and holding the values |  |  |
|  | 7. | ### Power Supply |  | Battery with low battery indication and 220V, 50Hz AC with adopter |  |  |
|  | 8. | ### Display |  | Back Lit LCD Display |  |  |
|  | 9. | ### Scope of Supply |  | Instruments with user manuals and calibration certificates. |  |  |
| - All the instruments shall be provided with protective carrying case. |  |  |  |  |  |  |

| BHEL | # TECHNICAL SPECIFICATION FOR digtal hand held instrumnets | Ref. : GT/SPEC/05/1014 Rev. : 00 Page 3 of 3 |
| --- | --- | --- |
| - The O&M manual shall consists of the following : - The Identification detail of equipment like BHEL PO no., Vendor’s job identification number, full contact address etc. - Brief Description of system. - System operation and maintenance requirement. - Operation and maintenance safety instruction. - Final data sheet and drawings - Test Reports and certificates - Recommended Spares. Inspection : - Calibration certificates and material certificates shall be furnished. - Inspection at manufacturing works by BHEL representative before despatch. Warranty : Instruments shall be have warranty of 18 months from date of commissioning or 24 months from date of supply. |  |  |