SPECIFICATION FOR  PORTABLE METALS ANALAYSER

Indent No. C/4200/64217/T

Date: 26.09.2006

- The Portable Metals Analyser shall be used for  the analysis of elemental constituents of metals & alloys, including ferrous and non ferrous metals, irrespective of sample size,shape & form.

- While the instrument is expected to operate in the grade analysis & pass-fail modes, the main usage of the instrument shall be accurate elemental analysis only. Supplier to ensure that the model quoted meets this primary requirement.

- The analyzer shall use a low power X-Ray tube for excitation of the elemental atoms,and, the analysis results shall be displayed either on the integrated screen or PDA.

- All elements from Titanium to Uranium in ferrous and non ferrous matrices shall be analysed & each analysis  shall be accompanied by percentage error of each element analysed.

- The matrices, the elements in these matrices, the range of concentration of the elements that are to be analysed in them are given in  Annex. I. For each of these matrices, the supplier shall provide data as to whether each of  the elements can be detected ,the detection range and, the percentage errors for this element within this range.

- The equipment as delivered should be capable of performing the analyses detailed above. If any matrix-element range is not in the standard protocol of supply/ display, the supplier should indicate the same along with the possibility of factory calibrating for these combinations.

- As far as the internal library of grades is concerned, the supplier shall give a list of grades with their compositions available in the library & the possibility of adding our plant grades in the library. Additionally, it is to be indicated as to how the grade determination is done i.e. the percentage deviation from the nominal composition that qualifies an analysis as belonging to a particular grade.

- The hardware specification of the equipment  shall meet the requirements given below. Supplier shall confirm that each of these parameters can be met.

    - Design		Ergonomic pistol grip design

    - X Ray Source	Low power X-Ray tube. The anode used shall be

such as to be capable of detecting  ***% Ag in Cu

matrix. Tube life min. 5 yrs guaranteed.

    - Detector		High resolution Peltier cooled Si-PIN type

    - Power Source	Rechargeable Li-ion batteries with hot swap

capability. There should be no data or standardization requirement during swap.Each battery to have minimum 6 hrs continuous operational capability.

    - Display & Control	Integrated touch screen or iPAQ with touch screen

Analyser should be controllable using this &

specimen I.D. etc. should be capable of being

entered using this.

    - Memory		At least 3000 test results with spectra

    - Standardization	Only during startup

    - Data Transfer	Data transfer to PC shall be possible. Hardware

& software for this to be supplied.

    - Standard accessories:

            - 2 Li-Ion battery packs

            - 220 V charger for battery pack

            - PC interface cable

            - PC data transfer software

            - Verification standard

            - Weld collimator

            - Benchtop instrument stand

- Along with the offer,the supplier shall give complete details of calibration /standardization procedure for the matrix & elemental ranges given above. Supplier to also clearly indicate as to whether a single standardization give accurate results all matrices,or, different types of standards are required,and,if so, what will be the standards required for the above range.

- Along with the main equipment, the supplier shall quote for the following also:

  - Spare X-Ray tube

  - Spare battery pack (Pl. indicate shelf life)

  - AC adapter for working directly off 220V, 50 Hz AC supply

  - Spare PDA (where applicable)

  - Adapter for wires & rods

  - Any other accessories/software available with detailed technical writeup regarding utility

  - Extra certified calibration standards with detailed compositions.

- The equipment shall be commissioned in our laboratory by the suppliers’ representative who shall demonstrate the complete functioning of the equipment.,and capability to meet our analytical requirements.

- The instrument shall be guaranteed for satisfactory working for 24 months from the date of commissioning.

- Along with the equipment, the supplier shall give 3 copies of operation & maintenance manuals with complete circuit diagram of the main equipment & all its accessories.

- Along with the  offer the supplier shall give:

  - The details of his service network, with details of service personnel & their experience in repair of the model of instrument offered,  especially in the northern region of the country.

  - List of users, with  name, complete address, e-mail etc.

- The offer should be accompanied by detailed technical literature . The offer should address each of the requirements given above with complete technical details. Compliance statements shall not be considered.

- Supplier to note that  the commitment against each & every clause above, shall be checked during commissioning & if not found complying the instrument is liable to be rejected.