# BHARAT HEAVY ELECTRICALS LIMITED					Date: ____________________

Heavy Power Equipment Plant, Ramachandra puram,			    Due Date: ____________________

Hyderabad-502 032. INDIA.						  Supplier Qtn. No.: ____________________

Phones: 23182250 / 23020173								Date: ____________________

# Purchase Dept.: 040 – 23020173 / 23021910 & 23021954

E-mail; burre@bhelhyd.co.in

## TENDER DOCUMENTS TO BE FILLED BY THE SUPPLIER FOR SUPPLY OF

SLAT SPRAY CHAMBER FOR LABS (M207CG7P0061/BRK)

| ### ADDRESS OF THE SUPPLIER: |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| TELEPHONE Nos: |  |  |  |  |  |  |
| FAX Nos.: |  |  |  |  |  |  |
| E-MAIL ADDRESS |  |  |  |  |  |  |
|  |  | Name of the Part | Description of BHEL Specification | Offered | Deviation | Remarks |
| 1.0 |  | Compliance |  |  |  |  |
|  | 1.1 |  | Equipment Shall be compliant to international standards. |  |  |  |
|  | 1.2 |  | Shall be Easy to install, use and practically maintenance free. |  |  |  |
| 2.0 |  | Chamber |  |  |  |  |
|  | 2.1 | Type | Equipment shall be cyclic corrosion test(CCT) chamber, i.e. it shall be used to expose the samples to variety of corrosion climates like condensation humidity, salt spray, and air drying |  |  |  |
|  | 2.2 | Capacity | Chamber capacity: Min of 450 liters. |  |  |  |
|  | 2.3 | Panels | Chamber capacity: min of 130 test panels of 6” x 4” |  |  |  |
|  | 2.4 | Reservoir capacity | Separate salt solution reservoir capacity of Min 100 liters with low level alarm shall be provided for ease of filling and monitoring the salt solution. |  |  |  |
|  | 2.4 | Distilled water | Distilled water reservoir capacity: Min 100 liters |  |  |  |
|  | 2.5 | Design | Chamber shall be designed to cycle between four conditions: Fog, Dry-Off, 100%Humidlty & Dwell. |  |  |  |
|  | 2.6 | Controls | Time & Temperature controlled by user friendly build in micro processor. |  |  |  |
| 3.0 |  | Temperature |  |  |  |  |
|  | 3.1 | #### Range | Cabinet temperature range: Min 350C |  |  |  |
|  | 3.2 | #### Accuracy | Temperature control accuracy: + or - 0.50C |  |  |  |
|  | 3.3 | #### Controlling system | Temperature controlling system: Solid state digital electronic (Temperature cum indicator with RTD sensor) |  |  |  |
|  | 3.4 | #### Protection | Over temperature protection shall be incorporated for both cabinet and as well as humidifier |  |  |  |
|  | 3.5 | #### Recorder | Cabinet temperature chart recorder shall be supplied with equipment. |  |  |  |
|  | 3.6 | #### Viewing window | Interior viewing window provision shall be provided for visual access to the cabinet interior samples under test. |  |  |  |
| 4.0 |  | #### Humidifier system |  |  |  |  |
|  | 4.1 |  | Humidifier system shall be compact internal air saturated system. Fully protected against over heating and low water. Automatic self filling system shall be available and fitted with overpressure relief valve. |  |  |  |
|  | 4.2 | #### Hour counter | Hour counter shall be incorporated |  |  |  |
|  | 4.3 | #### Automatic switch off | Equipment shall have automatic switch off facility after cyclic test |  |  |  |
|  | 4.4 | #### Air purging | Air purging facility shall be incorporated |  |  |  |
|  | 4.5 | #### Peristaltic pump | Peristaltic pump and graduated flow meter for precise dosing of salt solution |  |  |  |
| 5.0 |  | #### Pressure |  |  |  |  |
|  | 5.1 | #### Compressed air | Pressure of compressed air shall be up to 30 psi |  |  |  |
|  | 5.2 | #### Salt fog fall out rate | 1 – 2 ml solution / hr for 80 cm2 of horizontal collecting area. |  |  |  |
| 6.0 |  | #### General |  |  |  |  |
|  | 6.1 | #### Touch screen operator interface | Touch screen operator interface for ease of programming and monitoring of the test conditions. |  |  |  |
|  | 6.2 | #### Humidity sensor | Humidity sensor for display of the % of RH inside the chamber. |  |  |  |
|  | 6.3 | #### Pneumatic opening | Pneumatic opening and closing of roof for ease of operation. |  |  |  |
|  | 6.4 | #### On screen logging | On screen logging of cabinet temperature and humidity to enable the user to review what happened during any periods of unattended testing. |  |  |  |
| 7.0 |  | #### Power supply | Power supply: 230 V AC, 50Hz |  |  |  |
| 8.0 |  | #### Inspection | The supplier shall test the equipment as per the standard inspection and test procedure of the company and demonstrate the capabilities included in this specification in the presence of BHEL representatives before dispatch. |  |  |  |
| 9.0 |  | #### Experience | The supplier shall have a minimum experience of 5 years in manufacturing Salt Spray Chambers and should have supplied at least 10 equipment of the model quoted |  |  |  |
| 10.0 |  | #### No. of Installations | The supplier shall submit a users' list for the model quoted including the names of the user contact persons, telephone numbers and e mail addresses. |  |  |  |
| 11.0 |  | #### Training | Training on operation for a period of 1to 2 days for 2 persons shall be provided. |  |  |  |
| 12.0 |  | #### Service | The supplier shall explain in detail how the after sales service will be organized. The number of service engineers trained, qualified and experienced to attend service calls and details about ready availability of spares shall be mentioned. Time taken for service engineers to reach the site upon receipt of breakdown call shall be indicated and assured |  |  |  |
| 13.0 |  | #### Safety | The equipment shall fulfill safety regulation and guidelines |  |  |  |
| 14.0 |  | #### Commissioning | Installation and Commissioning shall be carried out at BHEL works |  |  |  |
| 15.0 |  | #### Spares | Spares for 2 years of trouble free operation shall be quoted. |  |  |  |
| 16.0 |  | #### Warranty | 18 months from date of commissioning |  |  |  |
| 17.0 |  | #### Manual | 3 sets of operation and maintenance manuals shall be provided |  |  |  |
| 18.0 |  | #### AMC | Supplier shall quote the AMC charges for 5 years after warranty period. |  |  |  |