## Prebid Queries
Client: | BHEL
End User: | Tata Kalinganagar
Item: | CV Analyzer (Gas Chromatograph)
RFQ No.: | T7I1S25943 (NIT_ 45253)
Sr. No. | Document No |  | Description | Bidder's Query | BHEL replies
| Clause No. | Page No.
1 | Annex. 1 | 2 | Supply of Calorific Value Analyzer System- 4 Nos. | We understand that total 4 nos. of GC to be considered; 2 nos. for COG & 2 nos. for mixed gas stream. All 4 nos. GC's are single stream GC. Please confirm our understanding is correct. | Confirmed
2 | Annex. 2, Fuel gas composition | 1 | Mixed Gas Composition | i) Kindly specify the component: CmHn in mixed gas composition ii) Provide the complete sample composition such that total is 100% iii) Provide the operating conditions (temperature, pressure) of mixed gas stream | i) Shall be provided after P.O. placement ii) Shall be provided after P.O. placement iii) Tentative parameters are: BFG+LDG: Line Size: 2620(OD)x10T Line MOC: API5LGr.B Pressure: 1200 to 1500mmWC Temp: 50 - 60 Deg C COG: Line Size: 1168(OD)x10T Line MOC: API5LGr.B Pressure: 900 to 1200mmWC Temp: 27 Deg C (amb) Exact details shall be provided after P.O. placement
3 | Doc.: PY56343; Cl. 8.7 (18) | 5 | CV analyzer shall have separate electronics and oven section with independent doors. | We understand that the Gas Chromatograph to be quoted shall be the latest model of a particular make with the basic design of i) separate electronics section & oven section with independent doors ii) electronic section on top & oven section in bottom Please confirm if our understanding is correct | Confirmed
4 | Doc.: PY56343; Cl. 8.7 (21 g) | 5 | Fuel Gas Data to Flow Meter via MODBUS | We understand that GC output is to be provided to DCS & Laptop only & not to any flowmeter OR flow computer. Please confirm. | Confirmed
5 | Doc.: PY56343; Cl. 13 | 6 | Mandatory Spares | Please note that Mandatory spares philosophy is not provided in the tender doc. So we understand that mandatory spares are not required. Please confirm. | Confirmed
6 | Annex. 3, GC Specs.; Cl. 13.1.1 (i) | 1 | The GCs are to be hooked up to the respective flow computers to communicate the analysis parameters to flow computer/s in order to have automatic calculation Compressibility Factor | We understand that, GC's are not required to communicate to flow computer & also compressibility factor calculation is not required. Please confirm. | Confirmed
7 | Annex. 3, GC Specs.; Cl. 13.1.1 (ii) | 1, 2 | The GC shall be installed at gas metering station for custody measurement purpose | We understand that GC is not required for custody measurement purpose & hence PTB certificate is also not required. Please confirm. | Confirmed
8 | Annex. 3, GC Specs.; Cl. 13.1.1 (ii, b) | 2 | Presently the requirement is for single stream analysis, however it should have extra stream expanding provision for taking independent sample from gas pipeline. | We understand that all 4 GC's are single stream only & no provision is required for future expansion. Please confirm. | Confirmed
9 | Annex. 3, GC Specs.; Cl. 13.1.1 (iv) | 2 | Based on component analysis, the GC will calculate the heating value of the sample gas and specific gravity as per latest version of ASTM D7314/ISO 7504 standard. | As mentioned in doc.: PY56343-R00-CV, GC shall calculate heating value & specific gravity as per ISO 6976. | Accepted; vendor may also follow ISO 7504
10 | Annex. 3, GC Specs.; Cl. 13.1.2 (I e) | 2 | Enclosure: IP65 or better | As per manufacturer standard, GC shall be IP54 which is mounted in field IP54 cabinet. Request for the acceptance. | Accepted. Bidder scope includes field cabinet also. Field Cabinet is independent for each GC
11 | Annex. 3, GC Specs.; Cl. 13.1.2 (ii m, n, o, q) | 3 | i) Repeatability Calorific value & Density: < 0.01% ii) Accuracy for Calorific Value & Density: < 0.1% iii) No. of streams for gas analysis: 3 (Three) iv) GC should have operating temp. upto 60 degC | Note that, i) Repeatability of GC shall be +/- 1 or 2 % FS for each measured component ii) Accuracy of measured component depends on accuracy of calibration blend iii) GC considered for single stream (not 3 stream) iv) Note that GC is mounted in field cabinet with vortex cooler. So 60 deg operating temp. is not required. | i) Bidder to provide installed base for further approval along with offer. Ii) Bidder to provide installed base for further approval along with offer. iii) Confirmed iv) Noted. It is installed in outdoor and hence, it shall operate at ambient condition if the cooler arrangement fails.