| SL. No. | PARTICULARS | REQUIREMENT |  |
| --- | --- | --- | --- |
| 1 | Purpose : | For measuring, recording and analysing the temperature of test object e.g. Turbogenerator Bar under High Voltage test. - During HV testing on Turbogenerator Bar, camera can be operated through remote controller or through computer. - The camera shall include rechargeable battery and LCD display. - The Thermal Imager shall be very user-friendly and lightweight and easy to carry. The equipment shall meet the following requirements / parameters : |  |
| 2. | SCOPE OF SUPPLY : |  |  |
| 2.1 | Main Equipment | The Thermo vision camera shall be with following technical parameters. |  |
| 2.1.1 | Image Performance | - Detector Type : Micro bolometer - Spectral band : ≈ (8 to 14 µm) - Frame Rate : 25 Hz. Min. - Field of View : 15° to 25° horizontal x 15° to 25° vertical - Focus Range : ≤ 500mm to infinity - Image zooming facility : Digital 1x to 8 x Continuous - Sensitivity : 0.1°K or better - Display type : 320 x 240 pixels or better |  |
| 2.1.2 | Measurement | Temperature measurement range : -5°C to 350°C or higher Accuracy : ± 2°C or ± 2% of reading or better |  |
| 2.1.3 | Operating Modes : | Black & white and colour Image Spot - movable for thermal analysis NOTE : Equipment should operate without liquid Nitrogen / or any other coolant. |  |
| 2.1.4 | Features : | The equipment should have the following features : |  |
|  | a ) | Simultaneous Visual and Thermal Imaging |  |
|  | b ) | Auto / motorized focus with remote operation facility |  |
|  | c ) | Auto Temperature Scan |  |
|  | d ) | Auto Temperature Range |  |
|  | e ) | View and external display on High resolution Colour LCD (size :3.5” minimum) |  |
|  | f ) | Image & Data storage |  |
|  | g ) | Image analysis w.r.t. temperature |  |
|  | h ) | Digital video Output with Real Time thermal image and Real Time video storage in PC through software for temperature analysis. |  |
|  | i ) | Extendable Memory facility may be indicated. |  |
|  | j ) | Facility for Emissivity Correction required. |  |
|  | k ) | Laser view finder for object identification |  |
| 2.1.5 | Operating Environment | Ambient Temperature : 0°C to 45°C Line Voltage & Frequency : 230V ± 10%, 50Hz ± 3 Hz. |  |
| 2.2 | Power Supply : |  |  |
|  | a ) | Rechargeable battery with Battery Charger from 230V AC mains with min. 2 Hrs. continuous operation. |  |
|  | b ) | Real time battery charging with auto cut-off facility. |  |
| 2.3 | Calibration : | Instrument Calibration Certificate shall be traceable to National / International Standard. The method of Calibration / agency to be indicated. |  |
| 2.4 | Commissioning & Training |  |  |
| 2.4.1 | Commissioning : | Party shall successfully commission the equipment at BHEL Haridwar works free of cost |  |
| 2.4.2 | Training : | Party shall provide training to 2 BHEL Engineers during commissioning of the equipment at BHEL Haridwar works free of cost |  |
| 2.5 | Other Requirements : |  |  |
| 2.5.1 | O & M Manual | 2 Nos. O & M Manual in original along-with 3 CD ROM’s of the equipment. |  |
| 2.5.2 | Test Certificate (TC) & Guarantee Certificate (GC) | Test Certificate (TC) & Guarantee Certificate (GC) of the equipment also to be supplied at the time of dispatch. |  |
| 2.5.3 | Packing List | Packing List shall be given with proper identification mark. |  |
| 2.5.4 | Packing | Party shall ensure for suitable packaging to avoid damage due to transportation. |  |
| 2.5.5 | Spares | Party shall recommend the necessary Spares ( if require ) separately of the main offer. |  |
| 2.5.6 | service facility | - Party shall also indicate in the offer regarding service facility in India - Annual Maintenance contract(AMC), if any, for the camera may also be indicated and separately be quoted.. |  |
| 3. | General Requirements : |  |  |
| 3.1 |  | Party may offer for different models meeting above requirements or better than that with essential and optional accessories and recommended spares. |  |
| 3.2 |  | Party shall offer the model with full catalogue. Any feature which is not covered in the catalogue as per our specification, party shall confirm those features in the offer separately. |  |
| 4. | Compliance Aspect | The offer shall contain confirmation for compliance to each and every clause of the specification. If there is any deviation, the same shall be clearly stated. |  |
| 5. | Offer Condition | Party shall submit the reference list with full address of customers to whom they have supplied the model worldwide and in India. |  |