Specification of   Portable  Vibration Analyzer:

- Item description: Portable Vibration Analyzer (FFT).

- Application:  Measurement  of    natural frequency of   Steam Turbine free standing, twisted  Low Pressure Moving Blades of lengths varying between 100 mm to 800 mm.  The accelerometer shall be attached to the twisted blade profile (with clay/magnet base of the  accelerometer) and the measurement shall be possible with single and gentle  strike on the blade profile.

- Number of channels: Minimum one.

- Frequency range: 5 Hz to 10000 Hz, linear scale in selectable ranges.

- Weight of the total system: Less than 10 kg.(including all accessories)

- Accuracy:  +/-  0.50% of FSD

- Power supply: Both Battery(rechargeable and 4 hrs back up time-minimum) and direct 230 volt power supply

- Ambient temperature: 5-45 Deg. Centigrade

- Weight of Accelerometer: ≤ 4 gms (Magnet base - optional)

- Display screen size: 125 mm (Length) minimum X 60 mm (Width) minimum.

- Display functionality: zoom, unzoom,   etc.

- Cursor for frequency and amplitude: to be provided.

- Resolution: 800 lines (minimum)

- Carrying case: capable to carry total instrument.

- Tuning hammer : 1 number (Optional)

- Internal memory: Capable to store 150 Spectrum data (minimum)

- Import/Export functionality: Capable to export Spectrum and frequency reading at the cursor for a spectrum into MS Excel format/other spreadsheet file to PC.

- Laptop of latest configuration to store spectrum and frequency data.

- Strong construction with a high resistance to adverse effects of water, dust, vibration and shock.

- Dynamic range: 60 db (minimum).

- Commissioning to be done at BHEL.

- Training for a  period of  six days in operation  &   maintenance  at  BHEL.

- The  supplier should have supplied  minimum of  three such instruments of  the same model.    Performance certificate   for satisfactory  working  over  past three years  for   minimum one  instrument   shall  be furnished.

- Spares and consumables  for  five  years’  trouble-free operation  shall be quoted.

- Warranty  for  two years shall be  provided.

(Jayanta Das)

DGM/QC/TC&GT