## Specs Annex-A
Bharat Heavy Electricals Limited, Heavy Electrical Equipment Plant, Haridwar - 249403, Uttarakhand, INDIA. |  |  | Enquiry No. & Date: | C6520/9/4167T/1
|  |  | Due Date: | 2009-09-11
Supplier's Ref.:
Date :
Technical Specification for Computed Radiography System
Note:-
1. The offer and all documents enclosed with offer should be in English language only.
2. The following format must be completely filled and deviations mentioned in the appropriate columns. Inadequate /incomplete, ambiguous, or unsustainable information against any of the clauses of the specifications / requirements shall be treated as non-compliance.
3. Vendor must submit complete information against clause no. 9 (Qualifying condition). The offer meeting this clause would only be processed.
NAME & ADDRESS OF THE SUPPLIER :
Telephone No.
Fax No.
e-mail :
Scope: Supply and Commissioning of Computed Radiography System with the following specifications:
Annexure 'A'
MATERIAL CODE NO.: ZFTH19402015
SCOPE: | SUPPLY AND COMMISSIONING OF COMPUTED RADIOGRAPHY SYSTEM WITH SPECIFICATIONS AS BELOW:
S. No. | Description - BHEL's Requirement | Specified | Offered | Deviations | Remarks
1.0.0 | Application
1.1.0 | The equipment will be mainly used for radiography of weld joints. Weld thickness varies from 1.6 mm to 30 mm. Few jobs are having thickness up to 50 mm. | Vendor to note
1.2.0 | Applicable Radiography specification is EN 1435 (1997) | Vendor to note
1.3.0 | Image Quality (Sensitivity) Requirements as per EN 1435 (1997), Annexure B | Vendor to note
1.4.0 | The equipment should be suitable for use with x-ray (upto 300 KV) and isotope (Ir-192 and Co-60) radiation sources. | Vendor to note
2.0.0 | Scope of Supply
2.1.0 | Supply and Commissioning of Computed Radiography System | Vendor to Comply
3.0.0 | Quantity Required
3.1.0 | 1 | 1 No.
4.0.0 | Major Technical Specification: Computed Radiography (CR) Scanner Requirements
4.1.0 | Conformance to Specifications and System Class:
4.1.1 | The scanner-imaging plate system class suitable for weld inspection (according to EN 14784-1 and ASTM E 2446) | Vendor to Specify
4.2.0 | Laser spot size (Maximum) | 50 μm
4.3.0 | Scanning Pixel Pitch / Pixel size (maximum) | 50 μm
4.4.0 | Bit Depth (Minimum) | 16 bit
4.5.0 | Main Equipment Weight (Maximum) | 30 Kg
4.6.0 | Device shall have an integrated eraser | Vendor to Comply
4.7.0 | Firewire (IEEE1394) or USB interface | Vendor to Comply
4.8.0 | Compatible carrying case shall be offered, if available. | Vendor to Specify and offer seperately
4.9.0 | Power Supply: 230 Volts (± 10%) AC, 50 Hz, Single phase | Vendor to Comply
4.10.0 | Satisfactory operation in ambient temperaturen upto 40º C | Vendor to Comply
5.0.0 | Consumables and Accessories
5.1.0 | Imaging Plates:
5.1.1 | Availability of high definition Imaging Plates suitable for weld inspection as per EN 1435. | Vendor to confirm
5.1.2 | The Imaging Plates should be resistant to scratches and damage during day to day handling. | Vendor to confirm
5.1.3 | Maximum Size of Imaging Plate required in our application | 30 cm X 38 cm
5.1.4 | Availability of Imaging Plates of sizes 10 cm X 19 cm, 10 cm X 30 cm, 10 cm X 38 cm, 15 cm X 19 cm and 30 cm X 38 cm | Vendor to confirm and offer each seperately
5.2.0 | Cassettes and other accessories:
5.2.1 | Availability of compatible cassettes with in built front and back screen | Vendor to confirm
5.2.2 | Availability of cassettes for Imaging Plates of sizes mentioned in clause 5.1.4 above | Vendor to confirm and offer each seperately
5.3.0 | Any other mandatory accessories (if any) required for the system operation shall be clearly mentioned. | Vendor to Specify and offer
5.4.0 | Optional accessories (if any) for the system. | Vendor to Specify and offer seperately
6.0.0 | Software
6.1.0 | Interfacing software to drive the equipment, if required, shall be offered. | Vendor to Specify
6.2.0 | Compatible software for image processing, if required, shall be offered. | Vendor to Specify
6.3.0 | Compatible software for digital image storage, if required, shall be offered. | Vendor to Specify
6.4.0 | Compatible software for defect evaluation, if available, shall be offered separately. | Vendor to Specify and offer seperately
6.5.0 | The software shall be able to exchange data / images with DICONDE (Digital Imaging and Communication in Nondestructive Evaluation) compliant equipments. | The vendor to confirm
7.0.0 | Imaging Workstation
7.1.0 | Imaging workstation requirement of compatible Computer / Laptop with suitable system configuration and high resolution screen. | Vendor to confirm and offer seperately
8.0.0 | Other Requirements:
8.1.0 | Test Certificate and Guarantee Certificate in original bearing the test date shall be submitted. | Vendor to Comply
8.2.0 | Certification from reputed certification bodies
8.2.1 | Mention certification from any reputed certification body [e.g. Federal Institute for Material Research and Testing (BAM), Germany etc.] , if any - for information only | Vendor to confirm
8.2.2 | In case answer is 'yes' for clause 8.2.1 , Copy of certificate from the certification body | Vendor to submit
8.3.0 | The supplier must give a guarantee for a period of at least two years from the date of commissioning of the equipment at BHEL, Hardwar. | Vendor to Comply
8.4.0 | Two copies of Operation and Maintenance Manual shall be supplied by the supplier. | Vendor to Comply
8.5.0 | Commissioning shall be done by the supplier at BHEL, Hardwar free of cost. Necessary training shall be provided by the supplier to NDT laboratory operating staff for at least five days. | Vendor to Comply
8.6.0 | For the imported machines Indian agents must attach the authority letter from the principal in the quotation stage itself. | Vendor to submit
8.7.0 | Maintenance spares required for five years running | Vendor to specify and offer seperately
8.8.0 | Suppliers will clearly mention the service facility available in India for the offered equipment. | Vendor to confirm
9.0.0 | Qualifying Conditions :
9.1.0 | Offer of only those vendors will be considered who have supplied and commissioned at least one similar type of computed radiography system in the previous five years, for industrial radiography in India or abroad. | Vendor to accept and confirm
9.2.0 | A list of customers to whom the similar computed radiography system has been supplied shall be provided mentioning the model of the equipment and year. | Vendor to inform
9.3.0 | Name / Designation of the customer’s contact person alongwith the address, contact phone / fax numbers, email address etc. shall be provided. | Vendor to inform
9.4.0 | BHEL reserves the right to verify information submitted by vendor. In case the information is found to be false / incorrect, the offer shall be rejected. | Vendor to accept & confirm