Specification of Thin Film Thickness Measuring Equipment

A. Application: Major application is for solar cell thin film thickness measurement

for films such as CIGS, CdS, ITO, ZnO etc. deposited on substrates such as glass,

metal foils, plastics. The equipment will also be used for measurement of thickness

of thin films of other materials.

B.  Main Features :

- Spectrometer, A / D converter, reflection probe and suitable software for measurement of thin films.

- Light source -  Suitable Tungsten halogen lamp

3.   Angle of incidence -   90˚ ( normal incidence )

4.   Effective wavelength covering a range of   400 nm to 1700 nm

5.   Type of detectors – Silicon array and InGaAs array (or any other  suitable)

covering 400 nm – 1700 nm wavelength range. Type of detector should be

mentioned in the offer.

6.   Thickness range of measurement -  Equipment should be capable of measuring

thickness of films in the thickness range of  20 nm – 100 μm.

7.   Number of layers -   The equipment should be capable of measuring single film

and multiple films (minimum 5 layers) deposited one over the other.

8.  The equipment should be capable of measuring thicknesses of typical layer

materials used in solar cell applications as mentioned below:

a) Copper-Indium-Gallium-Selenide(CIGS)

b) Cadmium Sulphide (CdS)

c) Zinc Oxide (ZnO)

d) Indium - Tin – Oxide (ITO)

Various other films that can be studied also should be mentioned.

9.   Supporting technical data: Graphical representation ( Reflectance ( % ) versus

wavelength (nm) of  incident light ) of actually measured samples using your

equipment for films  such as i) CIGS, ii) CdS, or combination should be enclosed

along with the thickness data..

If any such data for other materials is available it should be provided.

10  Optical fiber cable and probe -  Suitable optical fiber cable and probe for

transmitting light from source to the sample and to collect and re-transmit the

reflected light from the sample to the spectrometer should be provided.

11. Sample stand – Sample stand with probe holder suitable for measuring samples

at least up to 100 mm diameter should be provided. It should be capable of

accommodating substrates of thickness at least up to 10 mm.

12. Spot Size – A spot size within the range of  1 mm to 3 mm should be provided.

13.  Absolute accuracy -    < 1 % ( measured with SiO2 on Si )

14.  Reference material and thickness standard (SiO2 on Si wafer) for demonstration

should be supplied along with the equipment. Details to be provided.

15.  Thickness measurement precision -  0.1 nm

16.  Measurement of optical properties such as refractive index (n) and extinction

Coefficient (k) should be possible.

C.  Software :

1. Suitable software for measurement of thickness of single layer as well as of

multiple layers  (min. 5 different layers one above the other) to be provided.

2.  The software should be compatible with Microsoft Office products and it should

operate under Windows XP Environment.

3.  PC along with LCD monitor (17”) and printer to be provided with details.

All catalogues related to Software / Hardware to be provided. It is preferable to

supply PC available in India with software loaded and integrated with the system.

PC  of make Dell or HP or COMPACQ Desktop, with minimum configuration

Intel dual core CPU, 2.8 GHz, 2 GB DDR3, 320 GB HDD, DVD RW, 17” or

18.5” TFT Colour Monitor, Key Board, Optical Scroll Mouse, and HP PSC F 4488

Printer to be provided.

4.   Interface -   USB / RS232

5.  The computer and accessories should be Guaranteed for a period of 2 years from

the date of installation at R&D

D.  Power Input :

1.  The complete system comprising of spectrometer, computer, printer, UPS

etc. should be designed for use with 220 - 240 V A.C.(1 phase) with a mains

frequency of  50 ± 1% Hz.

2.  Suitable UPS (reputed make like APC/Leibert/SOCOMEC) with backup

power of up to 30 minutes to be quoted with details separately. The UPS should

be Guaranteed for a period of 2 years from the date of installation at R&D.

E.   Other aspects:

1.   The vendor should have supplied minimum two such equipments and the

addresses of the customers with telephone numbers should be furnished.

2.   Should have agents in India to provide after sales service and maintenance.

3.   The equipment should be Guaranteed for a period of 2 years from the date of

commissioning. During warranty period, if there is any repair to be carried out at

the suppliers works, transportation cost of equipment / component besides repair

/ replacement charges, if any, should be borne by the supplier.

4.    Catalogue related to each and every item should be enclosed.

5.    Dimensions of equipment, weight and space requirements to be given.

6.    Pre-installation requirements should be furnished.

7.    Installation & commissioning of the equipment to be carried out by supplier

at site at Corporate R&D, BHEL, Hyderabad, India. Training to be provided

at Corp. R&D, BHEL, Hyderabad.

8.    Commissioning charges, if any, to be indicated separately.

9.    Compliance statement of specification to be submitted along with the offer.

Without compliance statement, the offer is liable to be rejected. All tender

specifications  to be compared with equipment offered line by line and

documentary evidence must be enclosed by the supplier along with quotation.

10.   Two copies of operating manual to be provided along with the equipment.