Specification No.:QC-T/TKS/09/003                                                                                    Indent no :20092426

| Sr.No | Description | Required | Offered |
| --- | --- | --- | --- |
| 1.0 | Application: |  |  |
|  | The Portable vibration frequency measuring instrument is intended to be used in shop floor for natural frequency testing of turbine blades in static conditions. |  | Vendor to comply this column |
| 2.0 | Input: |  |  |
|  | No. of Channels | Two Nos. active |  |
|  | Connector Type: |  |  |
|  | For ICP type Accelerometer | BNC Type or 10-32UNF female |  |
|  | For Voltage Input | BNC Type |  |
|  | Dynamic range | >=120 dB |  |
|  | Coupling | AC/DC/ICP coupling |  |
| 3.0 | # Analysis Section: |  |  |
|  | Frequency Range | SUITABLE INSTRUMENT TO MEASURE FREQUENCY 20Hz TO 2000Hz atleast. |  |
|  | Frequency resolution | 6400 lines or more |  |
|  | Window Function | Hanning at least |  |
|  | Average Mode | Time (STA), Spectral, Linear Exponential peak hold and reference peak hold modes. |  |
|  | Measurement Mode | On-line and Off-line |  |
|  | Measuring accuracy | All input ranges at 1KHz:0.05dB (typical 0.015 dB) Temperature variability :< 0.1 dB/10 deg C |  |
|  | Analyser type | CHANNEL PROCESSING SHOULD BE REALTIME. |  |
| 4.0 | Displays: |  |  |
|  | Display Type | Time signal, Spectral, monitoring waterfall, Monitoring/ Compressed signal, Constant Band Tracking at least |  |
|  | Resolution | Input resolution: 24 bits (144dB) or better |  |
|  | Display Mode | Time signal, Spectral, Orbit, waterfall, Monitoring/ Compressed signal, Profile, Complex profile, Constant Band Tracking at least |  |
|  | Y/ Vertical axis unit | Any physical quantity in any physical unit (in Linear, Log and dB). Display in RMS, Peak, Peak-peak, RMS at least. |  |
|  | X/Horizontal axis | Linear, Logarithmic of frequency |  |
| 5.0 | Computer System (as per latest in market) |  |  |
|  | LAPTOP PC | Intel Dual Core 2 Duo Technology with latest processor speed, 4 GB DDR2 RAM, not less than 160 GB SATA Hard disk, optical scroll mouse, suitable display screen, MS office 2007, Adobe writer, three years warranty |  |
|  | Power & Network Cable | Power & Network Cable |  |
|  | Pen Drive | Pen Drive 4GB or More |  |
|  | Licensed installed operating system with software | Shall be based on Windows XP or latest one. |  |
|  | Licensed Installation CDs of Windows, software package for data acquisition & reporting and any other Softwares/Drivers. | Shall be provided as a Back-up. |  |
| 6.0 | General Specification: |  |  |
|  | Operating Temperature | 0 to 50 deg C |  |
|  | Relative humidity | 93% RH at 40 deg C non-condensing |  |
| 7.0 | Accessories: |  |  |
|  | ICP Type accelerometers | ICP Type with 10mV/g sensitivity, more than 7KHz frequency range, weight not more than 10gms, measuring range minimum 250 gpk, Connection 10-32 UNF or BNC at side or top mounting 10-32 UNF thread hole. |  |
|  | Cable (3 meter) of 2-3mm dia assembled with connector | Suitable flexible cable with both connector adjustable to accelerometer on one end and instrument at other end. |  |
|  | AC adapter (if required in instrument), Battery (Ni-Cd) is not preferable. | AC adaptor for running analyzer for AC supply. |  |
|  | Magnet base/ clip for clamping | Magnet base should be very light |  |
|  | Carrying case | It is preferable to offer the whole system(instrument, display LAPTOP) in one handy pack |  |
| 8.0 | General Conditions: |  |  |
|  | Vendors to comply that quotation covers all the items (i.e software, cables, Connectors, Adapters, magnet, accelerometer, computers etc.) even it is not specified in item list. The price quotation at 7.0 & 9.0 must be on single piece basis. |  |  |
|  | Vendors to comply that they shall commission the instruments at our works and build templates on each design of full set of blades for future testing. During commissioning Party should also demonstrate interfacing of instrument with personal computer & details of software. The party will provide service in future if any problem arises during operation. |  |  |
|  | The party will provide operation, service and trouble shooting manuals containing detailed circuit diagram. |  |  |
|  | The party will quote necessary spare parts / accessories items for minimum 05 years of satisfactory operation of Equipment even if not included at chapter 7.0 and 9.0 . |  |  |
|  | The party shall provide calibration certificate of instruments and accelerometer and explain the calibration procedures off instrument during commissioning. The party should quote for a contract of calibration in future. |  |  |
|  | The party can demonstrate the instrument at our works if required. |  |  |
|  | Guarantee period of instrument should not be less than 18 months counted from date of dispatch or 12 months counted from the date of commissioning. |  |  |
| 9.0 | Spares: |  |  |
|  | ICP Type accelerometers | Mentioned in accessories |  |
|  | Cables (3 meter) assembled with connector | Mentioned in accessories. |  |
|  | Battery if required |  |  |
|  | AC adapter |  |  |
|  | Magnet base/clip for clamping | Mentioned in accessories |  |
|  | Battery Charger: | If required in the instrument |  |
| 10.0 | Reporting software: |  |  |
|  | Final reporting should be automatically made on customized format (annexure-2) keeping in view of following details: |  |  |
|  | 1) During testing, earlier old data in supplied format Annexure-1 in excel programme should be activated and displayed in front of operator. |  |  |
|  | 2) old values under column PRESS would be available for comparison of data between PRESS and ROTOR |  |  |
|  | 3) Software should also be capable to generate data of same first two peaks into ‘ROTOR column” in annexure-1 during repeat testing of blades in future. In that case, operator shall move the cursor at “ROTOR” column for all rows of blade identification number. When testing of a particular blade is over, the value of first two peaks should be automatically entered into “ROTOR” column and differentiate value of PRESS and ROTOR should be calculated in “DIFF” column. |  |  |
|  | 4) The reporting tool has to be based on the analysis software. It has to be a specific development providing a macro which compares the current spectrum with a result template defending the acceptable natural frequencies of the turbine blade. The operator just need to set the blade, fix the accelerometer on it, run the macro and excite the blade using the hammer. The macro has to compare the acquisition spectrum with a result template. If the natural frequencies of the blade are acceptable, the results are automatically in the report (similar to Annexure-I). If the natural frequencies of the blades are not acceptable the operator is informed that the blade is faulty and no results is added to the report. Then the same process starts again with the second blade and so on. |  |  |
| 11.0 | Qualifying Conditions : |  |  |
|  | Only those vendors(OEMs or its authorized agencies), who have supplied and commissioned at least one number of instrument of same model within the past 10 years (from date of opening of tenders) and the referred instrument is presently working satisfactorily for more than one year after commissioning, should quote. However, if referred instrument has/ had been supplied to BHEL, then such instrument should be presently working satisfactorily for more than six months (from date of opening of tenders) after its commissioning and acceptance in BHEL. |  |  |
|  | The following information should be submitted by the vendor about the companies where referred instruments have been supplied. This is required from all vendors for qualification of their offer. |  |  |
|  | Name of the customer / company, where referred instrument(s) is(are) installed. |  |  |
|  | Complete postal address of the customer |  |  |
|  | Month & Year of Commissioning. |  |  |
|  | Specific application if any for which the instrument is supplied. |  |  |
|  | Name and designation of the contact person of the customer. |  |  |
|  | Phone , FAX no. and email address of the contact person of the customer. |  |  |
|  | Performance certificate from the customer regarding satisfactory performance of instruments supplied to them (original certificate or through e-mail directly from the customer). Original certificate may be returned after verification by BHEL, if required. |  |  |
|  | BHEL reserves the right to verify information submitted by vendor. In case the information is found to be false/incorrect, the offer shall be rejected. |  |  |