Indent No. CG7P0310

|  | BHARAT HEAVY ELECTRICAL LIMITED RAMACHANDRPURAM :: HYDERABAD 502 032. |  | Enquiry No. : |  |
| --- | --- | --- | --- | --- |
|  |  |  | Due Date : |  |
| HYDERABAD | CONTACT DETAILS |  | Supplier Qtn. No.: |  |
|  | DGM/PUR(CG) email : raogum@bhelhyd.co.in Phone No. 040-23182189 |  | Date : |  |
| TECHNICAL SPECIFICATION CUM COMPLIANCE CERTIFICATION FOR |  |  |  |  |
| PRECISION 6 VOLTAGE AND 6 CURRENT POWER ANALYSER |  |  |  |  |
| NOTE:- |  |  |  |  |
| 1. Vendor (OEM) must submit complete information against clause no. 7. The offer meeting this clause would only be processed. |  |  |  |  |
| 2. The "Offered" Column and where applicable, the "Deviations" & "Remarks" Column of this format shall be filled in by the Vendor and submitted along with the offer. Inadequate / incomplete, ambiguous, or unsustainable information against any of the clauses of the specifications/requirements shall be treated as non-compliance. |  |  |  |  |
| 3. The offer and all documents enclosed with offer should be in English language only. |  |  |  |  |
| 4.Vendor to submit offer in two parts in closed envelops viz. Technical bid and Price bid. |  |  |  |  |
| ADDRESS OF THE SUPPLIER : |  | ADDRESS OF THE INDIAN AGENTS : |  |  |
| TELEPHONE NOS.: |  | TELEPHONE NOS.: |  |  |
| FAX NOS.: |  | FAX NOS.: |  |  |
| E-MAIL ADDRESS : |  | E-MAIL ADDRESS : |  |  |
| SCOPE: DESIGN, MANUFACTURE, TESTING, CALIBRATION PACKING, DISPATCH, INSTALLATION, COMMISSIONING OF PRECISION 6 VOLTAGE AND 6 CURRENT POWER ANALYSER |  |  |  |  |

| 1.0.0 | Purpose of the Equipment: |  |  |  |
| --- | --- | --- | --- | --- |
| 1.1.0 | The AC/DC Power analyser is used for measurement and analysis of electrical parameters for Variable Frequency Drive power output(Motor option), Synchronous Generators, Brushless Exciters and Permanent Magnet Generators. |  |  |  |
| 2.0.0 | Technical specifications |  |  |  |
| 2.1.0 | The equipment shall be compact in size , Six elements (6-voltages and 6-currents ie 3-phase power input from two systems or six phase inputs for AC or 6 independent power inputs for DC power measurements) Display in numerical, waveform, phase angles, Trends in various formats. Wide current and voltage ranges . Display and print out of harmonics and various factors in numerical and graphical formats, data storage, GPIB, Ethernet ports for connecting instruments to controllers, data transfer from the storage. Analog /digital inputs and outputs for output of signals and external input of signals, User friendly key board and function keys. |  |  |  |
| 2.2.0 | No.of channels |  |  |  |
| 2.2.1 | Voltage : 6 AC or DC |  |  |  |
| 2.2.2 | Current : 6 AC or DC |  |  |  |
| 2.3.0 | Ranges |  |  |  |
| 2.3.1 | Voltage : 0-1000 V DC to 1 MHz 6 Channels ( 1 kV rms, 1.5 kV peak continuous), Basic Accuracy : 0.05% of rdg for DC and AC (Overall accuracy shall be better than 0.12 % at Full scale range including all the errors at 50/60 Hz) |  |  |  |
| 2.3.2 | Current : Three channels of 0-5 A DC to 1 MHz (or above) and external inputs from Shunt outputs [Selectable either direct current or through external inputs as voltage input] (7 A rms, 10 A peak continuous) External inputs peak of 5 times of range continuous Basic Accuracy : 0.05% of rdg for DC and AC (Overall accuracy shall be better than 0.12 % at Full scale range including all the errors at 50/60 Hz) (Provision for both Direct input of current and external inputs from shunt shall be available.) |  |  |  |
| 2.3.3 | Current : Three channels of 0-20 A DC to 1 MHz (or above) and external inputs from Shunt outputs [Selectable either direct current or through external inputs as voltage input] External inputs peak of 5 times of range continuous Basic Accuracy : 0.05% of rdg for DC and AC (Overall accuracy shall be better than 0.12 % at Full scale range including all the errors at 50/60 Hz) (Provision for both Direct input of current and external inputs from shunt shall be available.) Isolation for external shunt inputs shall be provided in the instrument. |  |  |  |
| 2.4.0 | Power measurement : The analyzer shall compute the power in all 6 power channels in AC and DC The overall accuracy of the power shall be 0.3% or better considering all the errors at 50/60 Hz. |  |  |  |
| 2.5.0 | Frequency measurements Upto 160 kHz. Accuracy 0.05% of reading |  |  |  |
| 2.6.0 | Sampling frequency : 100 kHz (Minimum) |  |  |  |
| 2.7.0 | Storage : The analyzer shall be provided with Data storage memory for measured values. |  |  |  |
| 2.8.0 | Measurement/calculation functions : Instrument shall be capable of measuring voltage and current and derive the following: Voltage, Current, Active power,  Active power, Reactive power, Reactive power, Apparent power, Apparent power, Power factor, Power factor, Phase angle, Phase angle, Current and voltage Peak measurements, Crest factor, Form factor, Impedance, resistance, Reactance, efficiency, Single phase 2 wire power (using one element) , Three phase three wire power( using two elements), Three phase four wire (using three elements), Harmonic measurement: Upto 99th harmonic of Voltage, current, power. Voltage THF,Current THF,Voltage TIF, Current TIF, Harmonic voltage factor(HVF), Harmonic current factor(HIF), THD of Voltage, current, power. Positive only voltages, Positive only currents, Positive only powers, Negative only voltages, Negative only currents, Negative only powers, rms, mean. ` |  |  |  |
| 2.9.0 | Display 5.7” Graphic display color or higher A variety of display formats for numerical values. ( measured and derived quantities) Display of minimum 3 waveforms Harmonic measurements, Phase diagram, Bar graphs for harmonic measurement, Trend display of measured / calculated values, Settings of instrument, external inputs, |  |  |  |
| 2.10.0 | Averaging functions : Normal,exponential,moving averages. |  |  |  |
| 2.11.0 | CT,PT ratio inputs : Provision shall be given for PT,CT ratio inputs. |  |  |  |
| 2.12.0 | Printer : inbuilt printer shall be provided or External printer connection shall be provided for Printing of screens, list of measured values, harmonic graphs, settings etc. |  |  |  |
| 2.13.0 | Interfaces and input/outputs |  |  |  |
| 2.13.1 | IEEE 488 (GPIB) and Ethernet(100 Base T) |  |  |  |
| 2.13.2 | Input for external triggering, pulse input for rpm, Analog input for torque. |  |  |  |
| 2.13.3 | Complete command reference for the interface available in instrument shall provided for customized programming. |  |  |  |
| 2.14.0 | Connecting cables |  |  |  |
| 2.14.1 | - Set of Connecting cables of 1.5 meter length including connectors at both the ends for measurement of each 6 voltages, 6 currents with direct connection - Set of cables for 6 shunt (current) connections including connectors at both the ends shall be provided. - The voltage rating of the cables shall be 2000V - The minimum current rating of 30 A rms for current cables. - Power supply, interface cables shall be provided. |  |  |  |
| 2.15.0 | Power Supply : 180 V to 260 V AC 50 Hz. |  |  |  |
| 2.16.0 | Operating temperature : 40 Deg.Cel. and storage temperature : 50 Deg.Cel |  |  |  |
| 3.0.0 | Testing and calibration |  |  |  |
| 3.1.0 | The complete supply shall be tested and calibrated at manufacturer works and reports shall be furnished. |  |  |  |
| 4.0.0 | Documents |  |  |  |
| 4.1.0 | The following documents shall be submitted along with the offer. - Compliance statement as per the BHEL specifications - Vendors Catalogs, literature and write ups - Basic drawings for the equipment offered their interconnectivity drawings. etc. - Reference list. - Extended specifications |  |  |  |
| 4.2.0 | The following documents shall be submitted along with the supply of equipment. - Test and calibration certificates for the equipment supplied. - Three sets of User manuals manuals - Three sets of maintenance manuals. |  |  |  |
| 5.0.0 | Guarantee : The supplied equipment shall be guaranteed for 18 months from the date of commissioning. |  |  |  |
| 6.0.0 | Qualification of vendors |  |  |  |
| 6.1.0 | Vendor should have supplied two such systems earlier. References of similar systems supplied earlier to be furnished with the offer. |  |  |  |
| 7.0.0 | Other conditions |  |  |  |
| 7.1.0 | The complete offered system or part of the system will be considered for procurement. |  |  |  |
| 7.2.0 | In case the quoting agency for this system is not OEM, then it is mandatory that the supplier should provide a letter from the respective OEM stating that, they will support the equipment in both technical and logistic for at least two years. |  |  |  |
| 7.3.0 | The system, any unit or parts there in, if found defective during commissioning or there after shall be replaced within a week of detection of such defects. |  |  |  |
| 8..0.0 | In case of any clarification required, the party taking part in tender shall contact BHEL and obtain written clarification regarding this tender. |  |  |  |

| Prepared by : | Checked and Approved by : |
| --- | --- |