## Sheet1
SCOPE: SUPPLY, INSTALLATION & COMMISSIONING FOR UPGRADATION OF EXISTING 50 KV ANALOG H.V. KIT TO 50 KV DIGITAL H.V. KIT WITH INPUT CURRENT LESS THAN 150 A BY USING LT REACTOR & INCREASE OF LEAKAGE CURRENT TO 10 A BY USING HT REACTOR
S. NO. | DESCRIPTION FOR BHEL REQUIRMENT | SPECIFIED / TO BE CONFIRMED BY | OFFERED | DEVIATIONS | REMARKS
1 | PURPOSE & WORKING MATERIAL
1.1 | Purpose :
| Existing 50 KV H.V. kit, model: T440048, manfacturing year 1963, AEI ENGLAND make, input current 610 A leakage current 5 A to be upgraded to input current less than 150 A by using LT reactor & increase of leakage current to 10 A by using HT reactor. | Vendor to note
1.2 | Job to be tested: High voltage testing of hydro generator machine up to 210 MW with measurement of leakage current, Tan-Delta testing | Vendor to note
2 | SPECIFICATION :
2.1 | All upgradation & modification is to be carried out in existing 50 KV HV Kit system (model TT440048, 1963, AEI make ) at BHEL Bhopal. Specification/features of existing HV kit system and its modified features are as mentioned below-
2.1.1 | Instrument should measure leakage current up to 10 A without tripping from its present 5 A leakage current rating . | Vendor to confirm
2.1.2 | The input supply side current reduce to less than 150 Amp from existing 610 A current rating | Vendor to confirm
2.1.3 | The operation of variac should be both manual & motorised. At present operation of variac is manual | Vendor to confirm
2.1.4 | All the indicating instruments should be digital with 0.21 accuracy. At present all metering are analog type | Vendor to confirm
2.1.5 | Overall accuracy of instruments after rectification should be 0.5 % or better. | Vendor to confirm
2.2 | Scope of Supply
2.2.1 | Control Panel for control & metering with following items- 1 Nos | Vendor to confirm
| Digital Meter for Output Voltage (kV) & Leakage Current (mA) | Vendor to confirm
| Digital Meter for input Voltage and Current | Vendor to confirm
| Timer, Flasher light , | Vendor to confirm
| Indication lamp & Push Button | Vendor to confirm
| Toggle Switches and selector switches | Vendor to confirm
| Control fuse and Trip circuit | Vendor to confirm
| Buzzer & Potentiometer | Vendor to confirm
| Limit Switches | Vendor to confirm
| Assembly, component fitment and wiring | Vendor to confirm
| MS sheet housing, powder coated with caster wheels | Vendor to confirm
| The Item Not Mentioned above but required for operation of HV kit shall be in the scope of vendor | Vendor to confirm
2.2.2 | LT Reactors 225 kVAR - ( 1 No. 150 kVAR + 1 No. 75 kVAR ) combined unit
| Input 415 V, 2-Phase, 50 Hz | Vendor to confirm
| Oil cooled type with first filling of oil, cooling natural type | Vendor to confirm
| Contactor panel for reactor selection (75 KVAR, 150 KVAR OR 225 KVAR) | Vendor to confirm
| Tank with heavy duty wheels | Vendor to confirm
2.2.3 | Supply part for Autotransformer (Variac)
| Geared Brake motor | Vendor to confirm
| Chain and sproket set | Vendor to confirm
| Maximum & Min limit switches | Vendor to confirm
| Zero limit switch | Vendor to confirm
| Contactor panel with motor starter | Vendor to confirm
| Machined drive shafts | Vendor to confirm
| Derachment type arrangement for auto and manual operation | Vendor to confirm
| Replacement of oil | Vendor to confirm
2.2.4 | HT Reactor 40 kV , 6.5 A Reactor
| Rating 260 kVAr , 40 kV fixed type | Vendor to confirm
| Oil cooled type with first filling of oil, cooling natural type | Vendor to confirm
| HT bushing 11 kV on earth side and 50 kV on HT side | Vendor to confirm
| Lifting Provision & Heavy duty wheels provision | Vendor to confirm
| Oil level guage, Breather, Oil drain Valve | Vendor to confirm
2.2.5 | Potential Transformer / Divider
| 50 kV PT or Divider | Vendor to confirm
| Burden 50 VA, Input 50 KV, output 220/110 V | Vendor to confirm
| Accuracy 0.5 % | Vendor to confirm
| Display Unit with connecting cable upto 5 Mtr | Vendor to confirm
| Cooling - Oil cooled | Vendor to confirm
2.3 | Any modification/components software and accessories not listed above required for complete system upgradation and successful operation of 50 kV HV kit system shall be in vendor's scope. | Vendor to confirm
2.5 | SAFETY ARRANGEMENTS:
| Following safety features in addition to other standard safety features should be provided to the instrument: | Vendor to confirm
| 1. Instrument should have adequate and reliable safety interlocks / devices to avoid damage to the system and workplace due to malfunctioning of equipment or operator mistakes. | Vendor to confirm
| 2. All electronics part are housed in shielded enclosures. | Vendor to confirm
| 3. During performing High Volatge test Electrical Tripping interlock shall be provided on primary side (LT side) as well as on secondry side (HT Side). | Vendor to confirm
2.6 | ENVORONMENTAL PERFORMANCE OF THE SYSTEM:
The instrument shall confirm to following factors related to environment:
| (a) If any safety / environmental protection encloser is required it should be built in the system by vendor. | Vendor to confirm
3 | ACCESSORIES:
3.1 | Vendor to furnish list and details of accessories (if any) to ensure smooth function/operation of HV kit system | Vendor to specify & confirm
4 | DOCUMENTATION : One set each of the following shall be supplied along with the system
4.1 | Instruction manual : Operating manual with maintenance features and troubleshooting Hard Copy - 2 Nos Soft Copy - 1 Nos | Vendor to confirm
4.2 | Test & Guarantee Certificate : One set | Vendor to confirm
4.3 | Calibration certificate to be submitted from laboratory: One set (Calibration certification of measuring intruments shall be submitted) | Vendor to confirm
5 | INSTALLATION & COMMISSIONING
5.1 | Supplier to take full responsibility for carrying out the installation and commissioning of the HV kit system. This includes installation of new material, interfacing with the existing system and demonstration of HV kit measurement for satisfactory working. Commissioning to be done in presence of technical expert from OEM. All the installation and commissioning work shall be done at BHEL Bhopal | Vendor to confirm
5.2 | Successful proving of supplied and installed components by the supplier shall be considered as part of commissioning. | Vendor to confirm
5.3 | Demonstration of the upgraded system to their full capacity. | Vendor to confirm
5.4 | Commissioning tools, required for commissioning of the system within stipulated time, shall be brought by the supplier on returnable basis. | Vendor to confirm
5.5 | Schedule for delivery, installation and Commissioning shall be submitted along with the offer. | Vendor to confirm
6 | CRITERIA FOR WORK COMPLETION
| Vendor to submit: | Vendor to confirm
*Test repots of all routine & operational tests carried out at supplier's works/BHEL works as per the specification.
*Calibration certificate from laboratory : One set
7 | INSPECTION
7.1 | Inspection of material as per scope of supply point no. 2.2.1, 2.2.2 and 2.2.4 shall be done at supplier works place. Supplier shall also perform the testing of LT and HT reactor at partial load at their work place.
8 | ACCEPTANCE CRITERIA:
8.1 | Successful interfacing with existing system and display parameters and function as per clause 2.1 to 2.3 at BHEL Bhopal | Vendor to confirm
9 | TRAINING:
9.1 | Training to BHEL, testing engineers and maintenance person should be provided at BHEL Bhopal in the area of - | Vendor to specify & confirm
(a) Use of all system Features, programing for Measuring System & supplied accessories etc.
(b) Electrical, Electronic & maintenance for system & other supplied equipments
10 | AMBIENT CONDITIONS & THERMAL STABILITY :
10.1 | Total System and all supplied items should work trouble free and efficiently under following condition and specified accuracies. | Vendor to confirm
Power supply :
Voltage : 415 V ± 10 %
Frequency : 50 Hz ± 1.5 Hz
Operating Humidity : 10% to 80 %
0perating Temperature : 5 °C to 45 °C
11 | PACKING :
| Air worthy/sea worthy & rigid packing for all items of complete system, all accessories and other supplied items to avoid any damage/loss in transit. When system is dispatched in containers, all small loose items shall be suitably packed in boxes | Vendor to confirm
12 | GUARANTEE :
12.1 | 12 Months from the date of commissioning of the system for new supplied components | Vendor to confirm
13 | GENERAL : The vendor should submit the following information:
13.1 | Detailed catalogues, drawinges/photographs of the system and accessories/ attachments should be submitted with the offer. | Vendor to confirm and submit
13.2 | Vendor to submit detailed scope of supply along with the bid. | Vendor to confirm and furnish
13.3 | Vendor to give point wise confirmation of this specification without which the offer is liable to be rejected. | Vendor to confirm and furnish
13.4 | Vendor can visit BHEL bhopal works before technical bid opening for study and estimate the work required for upgradation for HV kit system | Vendor to note
14 | QUALIFYING CONDITIONS :
14.1 | The supplier should either be an Original equipment manufacturer (OEM) or an authorised dealer of the OEM for the offered equipment. Dealers have to submit a valid certificate of authorisation from OEM for quoting the equipment, alongwith declaration for support from OEM for commissioning, after sales service & necessary spares. | Vendor to specify & submit
14.2 | The OEM or its dealer must have supplied and commissioned at least one equipment ( model may be different however meeting minimum requirements of 50 KV HV kit along with cl 2.2.2 & 2.2.4) of OEM make only in the past seven years (on the date of opening of Tender) | Vendor to confirm
14.3 | The supplier shall submit one PO copy in name of OEM or any dealer/vendor for equipment supplied in line with clause 14.2 along with Performance certificate of same PO (issued within 3 years of tender opening date) from the costomers regarding satisfactory performance of the equipment supplied to them. The performance certificate shall be minimum one year. (BHEL reserves the right to verify information submitted by vendor. In case the information is found to be false/incorrect, the offer shall be rejected.) | Vendor to submit
15 | POINTWISE CONFIRMATION
15.1 | Vendor should confirm / calrify pointwise (all the point) as per specification and provide original technical leaflet, technical details, photographs, scope of supply etc. at the first instance. | Vendor to confirm & Submit