Technical specification

Name of the equipment: Electronic height gauge

| Sl no | Description of BHEL requirement |
| --- | --- |
| 1 | Purpose: The equipment shall be capable of measuring accurate linear dimensions and suitable for operating in normal shop-floor conditions. |
| 2 | The equipment shall have the following features: |
| 2.1 | Measuring range : 0 to 800 mm or higher |
| 2.2 | Stroke : 0 to 600 mm or more |
| 2.3 | Resolution (selectable): 0.01 and 0.001 mm |
| 2.4 | Maximum permissible error : 2 +L/400 microns or better |
| 2.5 | Repeatability : Within 2 microns or better |
| 2.6 | Maximum squareness deviation (both frontal and lateral) : 15 microns |
| 2.7 | Display :- Digital |
| 2.8 | Power supply : (a) Both rechargeable battery and 230 VAC/50 Hz mains (b) 1 No. Battery Charger shall be supplied. |
| 2.9 | Operating temperature : 5 to 40 deg C |
| 2.10 | Built in air pump for easy movement of the Height gauge over Granite Surface Plate shall be provided. |
| 3 | The instrument shall be able to measure following parameters. - Height Measurement. - Inside/Outside Diameter Measurement. - Centre Height Measurement. - Centre-to-Centre Height Measurement. - Zero setting at any position. - Input of preset value - Perpendicularity measurement |
| 4 | Following probes shall be supplied: (a)Spherical Probes of diameter 4 mm - 2 Nos. (b)Spherical Probes of diameter 10 mm – 2 Nos. (c) Cylindrical probe – dia. 3 mm – 2 Nos. (d) Probe for measuring perpendicularity including its accessories - 1 set. |
| 5 | Following accessories shall be supplied: (a) Accessory for measuring hole depth (upto 20 mm depth ) in vertical direction (b) Probe setting piece. |
| 6 | Spares for 5 years for trouble free operation with itemized break-up shall be quoted. |
| 7 | Commissioning and demonstration of trouble-free operation of the equipment shall be done at BHEL |
| 8 | Training in operation & maintenance shall be provided at BHEL. |
| 9 | Operation and maintenance manual in English shall be supplied. |
| 10 | Calibration certificate for the equipment / accessories shall be supplied. |