# Technical Specifications

Film Processing Tank and Illuminator

(Indent No.20084527, Dated 27.10.2008)

| Sl.No. | #### Technical Parameters | # BHEL Specification # Requirement | Relevant technical detail of the offered equipment shall be filled against each row by the bidder. |
| --- | --- | --- | --- |
|  | Film Processing Tank (1 No.) |  |  |
| 1. | Overall dimensions | length 150 cm. x width 60 cm. (including valves / plumbing in the back) x height 90 cm. |  |
| 2. | Minimum ground clearance | 6 inches. |  |
| 3. | Number of processing tanks | Developer, stop bath and fixer tank suspended in one master tank filled with water. One running water tank and a separate wetting agent tank. |  |
| 4. | Material of tanks | Austenitic stainless steel to avoid rusting. |  |
| 5. | #### Provision for cooling and heating | Provision for automatic cooling and heating of master tank to maintain 20°C temperature. |  |
| 6. | #### Temperature control for master tank | Automatic temperature control through thermostat. |  |
| 7. | #### Location of cooling unit | Cooling unit on the right side and tanks on the left side. |  |
| 8. | Width and capacity of developer / fixer tanks | Width 12.5 cm. (minimum) and capacity 22.5 litres (minimum). |  |
| 9. | Control panel | - Main electric power on / off switch alongwith MCB. - Digital temperature indicator (in °C) for master tank. - Electric control switches of the water control valves to be provided on control panel. - Indicating lights for power on, cooling on, heating on to be provided. |  |
| 10. | Insulation | Good insulation around master tank to minimize the heat gain from the atmosphere during summer season and heat loss to the surrounding during winter season. |  |
| 11. | #### Plumbing | Water inlet and water outlet from the left side of the tank. |  |
| 12. | Valves | Electrically operated valves to fill the master tank and to drain the master tank and running water tank. |  |
| 13. | Power Supply | 230 Volts, 50 Hz, single phase AC. |  |
| 14. | Ambient temperature | Upto 45°C in summer. |  |
| 15. | Operation/ maintenance manual | At least 3 copies to be provided. |  |
| 16. | #### Guarantee | At least 12 months guarantee against manufacturing defects including cooling unit from the date of commissioning. |  |
| 17. | Commissioning / training | Equipment commissioning by the supplier at B.H.E.L. works. |  |
|  | Illuminator (1 No.) |  |  |
| 1. | Illuminating capacity | Capable of illuminating upto optical density 4.0 |  |
| 2. | Viewing window size | 12 inches X 4 inches (Minimum), adjustable viewing window. |  |
| 3. | Illumination control | Continuous variation illumination control knob. |  |
| 4. | Cooling | Adequate arrangement for dissipation of heat generated by illuminator main lamp(s). |  |
| 5. | Background light | A small bulb that lights when main illuminator light is off. |  |
| 6. | #### Foot padel | Foot padel controlled illuminator lighting. |  |
| 7. | Power Supply | 230 Volts, 50 Hz, single phase AC. |  |
| 8. | Ambient temperature | Upto 45°C in summer. |  |
| 9. | #### Guarantee | At least 12 months guarantee against manufacturing defects to be provided. |  |