| fnukad ,oa gLrk{kj SIGN & DATE |  |  |  | PRODUCT PURCHASE SPECIFICATION |  |  |  |  | # Capital |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |  | ## Ik`"B |  |  | ## 1 | ## dk |  | ## 3 |
|  |  |  |  |  |  |  |  |  | ## Page |  |  | 1 | of |  | 3 |
| SUPERSEDES INVENTORY NO. |  | Multi-Gas Monitor - CO, H2S, Cl2 & O2 - GENERAL: Equipment is required to measure the various gases in work environment of Industry. - SPECIFICATION OF EQUIPMENT: Multi Gas Monitor suitable for measuring concentration of gases in work place and process emissions on enclosed areas. - Parameters to be measured : Lower Explosive Limit (LEL) as Methane, Oxygen, Hydrogen Sulphide, Carbon Monoxide, Sulphur Dioxide, Nitric Oxide, Nitrogen Dioxide, Chlorine and Ammonia. Multi Gas Monitor includes :- - Carrying case - Filters - Alkaline battery adapter. - Rechargeable Nickel Metal Hydride battery & charger. - Software for data logging suited for windows. - Computer interface cable. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| lkexzh lwph la[;k dks vf/kdzfer djrk gS |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| COPYRIGHT AND CONFIDENTIAL The information on this documents is the property of Bharat Heavy Electrical Limited.It must not be used directly or indirectly in any way detrimental to the interest of the company |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| LoRokf/kdkj ,oa xksiuh; bl izys[k esa nh xbZ lwpuk Hkkjr gsoh bysfDVªdYl dh lEifRr gS bldk izR;{k ,oa vizR;{k :Ik ls fdlh Hkh rjg iz;ksx] tks fd daiuh ds fgr esa gkfudkjd gks u fd;k tk, A |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| gLrk{kj ,oa fnukad SIGN & DATE |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  | uke NAME |  | fnukad ,oa gLrk{kj SIGNATURE & DATE |  |  |  |  |  |
|  |  |  | A. Goel |  |  | vuqoknd TRANSLATED BY |  | Bhoj raj |  |  |  |  |  |  |  |
|  |  |  |  |  |  | fuekZ.kdrkZ WORKED BY |  | Bhoj Raj |  |  |  |  |  |  |  |
| Lkkexzh lwph la[;k INVENTORY NO. |  |  |  |  |  | tkapdrkZ CHECKED BY |  | A. Goel |  |  |  |  |  |  |  |
|  |  | Lger foHkkx AGREED DEPTT. | uke NAME |  | fnukad ,oa gLrk{kj DATE & SIGNATURE | i;Zos{k.kdrkZ SUPERVISED BY |  | A. Goel |  |  |  |  |  |  |  |
|  |  |  |  |  |  | Lohd`fr % APPROVED : |  |  |  |  |  |  |  |  |  |
|  |  | REV.NO. 1 |  |  |  | fuekZ.k % PREPARED : | T+kkjh % ISSUED : |  |  |  | fnukad % DATE : |  |  |  |  |
|  |  | Dt. 1.11.08 |  |  |  |  |  |  |  |  |  |  |  |  |  |

| *+- fnukad ,oa gLrk{kj SIGN & DATE |  |  |  | PRODUCT PURCHASE SPECIFICATION |  |  | # Capital |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | ## Ik`"B | ## 2 |  | ## dk | ## 3 |
|  |  |  |  |  |  |  | ## Page | 2 |  | of | 3 |
| SUPERSEDES INVENTORY NO. |  | SENSORS REQUIRED: 1. Methane Sensor (Range: 1 - 100% Vol./LEL, Resolution: ± 1% ) 2. Oxygen Sensor (Range: 0 - 25 % Vol., Resolution: ± 0.25%) 3. Hydrogen Sulfide Sensor ( Range: 0 - 100 ppm, Resolution: ± 1 ppm ) 4. Carbon Monoxide Sensor ( Range: 0 - 500 ppm, Resolution: ± 1 ppm ) 5. Chlorine Sensor ( Range: 0 - 10 ppm, Resolution:± 0.1ppm) 6. Ammonia Sensor ( Range: 0 - 50 ppm, Resolution: ± 1 ppm ) 7. Nitric Oxide Sensor ( Range: 0 - 250 ppm, Resolution: ± 1 ppm ) 8. Nitrogen Dioxide Sensor ( Range: 0 - 20 ppm , Resolution: ± 0.1ppm) 9. Sulphur Dioxide Sensor ( Range: 0 - 20 ppm, Resolution: ± 0.1ppm) 10. Calibration gas : Portable Cylinder of Mixture of Gases or known concentration The Carbon Monoxide should be based on Infra Red spectroscopy. Others sensors should be based on electrochemical. 3.0 General Terms & Conditions 3.1 Installation & Commissioning : Instrument should be commissioned at PCRI, BHEL Haridwar 3.2 Document : -One hard copies O&M manual in English language giving complete technical information for operation. - Calibration certificate to be provided from recognized agency / laboratory. 3.3 Guarantee: The equipment shall be guaranteed for any breakage / malfunctioning of the apparatus for a period of 18 months from the date of commissioning or 24 months from the date of dispatch, whichever is earlier. In case of any breakdown during this time period, the guarantee period will get extended accordingly. TC & GC certificate to be provided. |  |  |  |  |  |  |  |  |  |
| lkexzh lwph la[;k dks vf/kdzfer djrk gS |  |  |  |  |  |  |  |  |  |  |  |
| COPYRIGHT AND CONFIDENTIAL The information on this documents is the property of Bharat Heavy Electrical Limited.It must not be used directly or indirectly in any way detrimental to the interest of the company |  |  |  |  |  |  |  |  |  |  |  |
| LoRokf/kdkj ,oa xksiuh; bl izys[k esa nh xbZ lwpuk Hkkjr gsoh bysfDVªdYl dh lEifRr gS bldk izR;{k ,oa vizR;{k :Ik ls fdlh Hkh rjg iz;ksx] tks fd daiuh ds fgr esa gkfudkjd gks u fd;k tk, A |  |  |  |  |  |  |  |  |  |  |  |
| gLrk{kj ,oa fnukad SIGN & DATE |  |  |  |  |  |  |  |  |  |  |  |
| Lkkexzh lwph la[;k INVENTORY NO. |  | REV.NO. 1 |  |  | fuekZ.kdrkZ WORKED BY | Bhoj Raj |  |  |  |  |  |
|  |  |  |  |  | TkkapdrkZ CHECKED BY | A. Goel |  |  |  |  |  |
| *+- fnukad ,oa gLrk{kj SIGN & DATE |  |  |  | PRODUCT PURCHASE SPECIFICATION |  |  | # PR00215 |  |  |  |  |
|  |  |  |  |  |  |  | ## Ik`"B | ## 3 |  | ## dk | ## 3 |
|  |  |  |  |  |  |  | ## Page | 3 |  | of | 3 |
| SUPERSEDES INVENTORY NO. |  | 3.4 Users : Along with the offer, the supplier shall give a list of users of similar equipment in India. It should contain detailed name, address & department, telephone number, e-mail address, model, year of supply etc. 3.5 Packing & Dispatch: The equipment / accessories shall be packed in a strong suitable water proof / vibration proof packing boxes capable of bearing road transit hazards. Packing box shall be properly identified and marked with BHEL's(India) Purchase Order Number. 3.6 All tools, tackles, gases, commissioning spares, consumables, calibration standards to suit Indian conditions, maintenance tool kit necessary for installation & commissioning & calibration of equipment at site be included in the offer. 4.0. Cross referred Document: Nil |  |  |  |  |  |  |  |  |  |
| lkexzh lwph la[;k dks vf/kdzfer djrk gS |  |  |  |  |  |  |  |  |  |  |  |
| COPYRIGHT AND CONFIDENTIAL The information on this documents is the property of Bharat Heavy Electrical Limited.It must not be used directly or indirectly in any way detrimental to the interest of the company |  |  |  |  |  |  |  |  |  |  |  |
| LoRokf/kdkj ,oa xksiuh; bl izys[k esa nh xbZ lwpuk Hkkjr gsoh bysfDVªdYl dh lEifRr gS bldk izR;{k ,oa vizR;{k :Ik ls fdlh Hkh rjg iz;ksx] tks fd daiuh ds fgr esa gkfudkjd gks u fd;k tk, A |  |  |  |  |  |  |  |  |  |  |  |
| gLrk{kj ,oa fnukad SIGN & DATE |  |  |  |  |  |  |  |  |  |  |  |
| Lkkexzh lwph la[;k INVENTORY NO. |  | REV.NO. 1 |  |  | fuekZ.kdrkZ WORKED BY | Bhoj Raj |  |  |  |  |  |
|  |  |  |  |  | TkkapdrkZ CHECKED BY | A. Goel |  |  |  |  |  |