|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 Rev 01 page 1 of 5 |
| --- | --- | --- |
| 1.0 TECHNICAL SPECIFICATIONS OF STANDARD PRESSURE GAUGES FOR TURBINE OIL GRADE ISO VG-46 - Type : Bourdon - Medium : Turbine oil grade ISO VG-46 - Working Temperature : 70 Deg C - Accuracy : ±1% of FSD - Mounting : Direct with bottom entry - Pressure Range : as per scope of supply in SECTION-I - Dial Type : Black figure with white background with anti-corrosive finished scale 9. Window : Acrylic/ Laminated safety Glass 10.Enclosure : SS/ die- cast aluminium housing dust proof and splash proof to IP-65 11.Process connection : ¼” NPT 12.Movement : SS-304 Each pressure gauge should be calibrated for Fluid turbine oil grade ISO VG-46. |  |  |

|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 Rev 01 page 2 of 5 |
| --- | --- | --- |
| 2.0 TECHNICAL SPECIFICATIONS OF STANDARD PRESSURE GAUGES FOR WATER - Type : Bourdon - Medium : Distilled water - Working Temperature : 70 Deg C - Accuracy : ±1% of FSD - Mounting : Direct with bottom entry - Pressure Range : as per scope of supply in SECTION-I - Dial Type : Black figure with white background with anti-corrosive finished scale 9. Window : Acrylic/ Laminated safety Glass 10.Enclosure : SS/ die- cast aluminium housing dust proof and splash proof to IP-65 11.Process connection : ¼” NPT 12.Movement : SS-304 Each pressure gauge should be calibrated for Water. |  |  |

|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 Rev 01 page 3 of 5 |
| --- | --- | --- |
| 3.0 TECHNICAL SPECIFICATIONS OF DIFFERENTIAL PRESSURE GAUGES FOR TURBINE OIL GRADE ISO VG-46 1.Type : Bourdon 2.Medium :Turbine oil grade ISO VG-46 3.Working Temperature : 70 Deg C 4.Accuracy : ±0.1% of FSD 5.Mounting : Direct with bottom entry 6.Pressure Range : as per scope of supply in SECTION-I 7.Dial Type : Black figure with white background with anti-corrosive finished scale 8. Window : Acrylic/ Laminated safety Glass 9.Enclosure : SS/ die- cast aluminium housing dust proof and splash proof to IP-65 10.Process connection : ¼” NPT 11.Movement : SS-304 12.Sensing Element : Diaphragm 13.Diaphragm material : SS-316 14.Diaphragm Chamber : SS-304 15. Max. Stat Pressure : as per scope of supply Each differential pressure gauge should be calibrated for Fluid turbine oil grade ISO VG-46. |  |  |
|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 Rev 01 page 4 of 5 |
| 4.0 TECHNICAL SPECIFICATIONS OF PRECISION PRESSURE GAUGE INDICATOR - Type : Bourdon - Medium : Air & Hydrogen - Working Temperature : 70 Deg C - Accuracy : ±0.1 % of FSD or better - Pressure Range : as per scope of supply in SECTION-I - Dial Type : Black figure with white background with anti-corrosive finished scale 7. Window : Acrylic/ Laminated safety Glass 8. Enclosure : SS/ die- cast aluminium housing dust proof and splash proof to IP-65 Each pressure gauge should be calibrated for Air & Hydrogen. |  |  |

|  | SPECIFICATION FOR PRESSURE GAUGES REQUIRED AT TEST BED,BLOCK-1 | SPECIFICATION NO. TB-C02-0809 Rev 01 page 5 of 5 |
| --- | --- | --- |
| SECTION-I SCOPE OF SUPPLY The Scope of Supply is as follows: Description; Range; Dial Size (in mm); Least count (in kg/cm²); Qty.(nos.). STANDARD PRESSURE GAUGES FOR TURBINE OIL ISO VG-46; 0-6kg/cm²; 150; 0.1; 20. _________ ,,__________; 0-10kg/cm²; 150; 0.1; 20. _________ ,,__________; 0-250kg/cm²; 100; 10; 06. STANDARD PRESSURE GAUGES FOR WATER; 0-16kg/cm²; 150; 0.5; 06. DIFFERENTIAL PRESSURE GAUGES FOR TURBINE OIL ISO VG-46 i) maximum static pressure suitable for 10 kg/cm2 oil line (party to mention in offer) ii) maximum static pressure suitable for 250 kg/cm2 oil line (party to mention in offer); 0-2kg/cm²; 150; 0.02; 03 01. PRECISION PRESSURE GAUGE INDICATOR; 0-7kg/cm²; 250; 0.01; 02. |