Scope of Supply for Instrumentation required for commissioning and field-testing of the equipment

- Multimeter

Multimeter shall meet following technical requirements:-

- Multimeter shall be 6,000 Count Resolution with true RMS voltage and current measurements capability.

- Multimeter shall have 0.09% dc accuracy, measure current up to 10A ac/dc, measure voltage up to 1000V ac and dc , frequency  up to 100 KHz, capacitance up to 10000µf, resistance 50 M Ohm, continuity and diode check facility.

- Multimeter shall a digital LCD display with an analog bar graph.

- Multimeter shall have frequency, capacitance, resistance, continuity and diode measurements.

- Multimeter shall be independently tested for safe use in CATIV 600V/ CAT III 1000V environments. All inputs are protected to IEN61010-1 CATIV 600V/CAT III 1000V.

- Calibration Certificate valid for 1 year.

- Warranty: Minimum 1 years from date of supply.

- Clamp Meter

Clamp meter shall meet following technical requirements

- AC current measurement range: Up to 1000 Amp with resolution less than 1 Amp, Accuracy 2% or better.

- DC current measurement range: Up to 1000 Amp with resolution less than 1 Amp, Accuracy 2% or better.

- AC voltage measurement range: Up to 1000 V with resolution less than 1 V, Accuracy 2% or better.

- DC voltage measurement range: Up to 1000 V with resolution less than 1 V, Accuracy 2% or better.

- Resistance measurement range: Up to 60 KΩ with accuracy 1% or better

- Capacitance measurement range: Up to 1 mF with accuracy 1% or better

- Frequency measurement (AC coupling) range: Up to 1 kHz with accuracy less than 1%.

- Measurement category:  CAT III 1000V, CAT IV 600 V

- Jaw opening:  Up to 40 mm or higher

- Operating temp: 0°C - 50°C

- Applicable Standard: EMC-EN61326, Safety-IEC 61010

- Accessories included: Battery, Test leads with one additional set.

- Calibration Certificate valid for 1 year

- Warranty: Minimum 1 years from date of supply

- Vane anemometer

Vane Anemometer shall meet following technical requirements:-

Measuring Range: 0.4 to 20 m/s
Accuracy: ±0.2 with resolution 0.1m/s
Operating Temperature: -10°C to 50°C
Data Hold – Yes
Automatic Power Off – Yes

Warranty: - Minimum 1 years from date of supply