Syringe Pump

Specification:

- INFUSION FLOWRATERANGE0.1-1200 ml/h (depending on the syringe capacity).

- FLOW RATE ACCURACY±1% on mechanism ; ±2 % on syringes.

- SYRINGE CAPACITIES  5,10, 20, 30/35, 50/60 CC.

- VOLUME/DOSE INFUSEDVolume: 0,1-999,9 ml; Dose: starts at 0.1 ml.

- INFUSION MODES  ml/h.

•Dose rate.

Dose units: ng, µg, mg, cal, kcal, U, kU, mmol, mol, Kmol / kg.

Time units: 24h / h / min.

•Volume or dose / time: 0.1-99.9 units ; 00h01- 96h00.

• Volume limit: 0.1-999.9 ml.

- INFUSED Volume: 0,1-999,9 ml; Dose: starts at 0.1 ml.

- BOLUS Manual bolus: Rate: 50 ml/h 50 -1200 ml/h; 50 ml/h.

- Programmed bolus (dose or volume/time): 0.1-99.9 units / 0.1-1200 ml time auto-calculation.

Drug labelling: infusion with the drug names

PAUSE Programmable from 1 minute to 24       hours, increments from minute to minute.

- DRUG LABELLING Display of the drug name based on a list of  pre-set drug names.

PRESSURE  MANAGEMENT  PRESSURE MODES  2 modes available: variable or 3 pre-set levels - From 100 to 900 mmHg, 50 mmHg increment.

DPS The Dynamic Pressure System - DPS - warns of pressure variations. A risk of obstruction or a possible leak in the infusion line can thus be anticipated.

- PRESSURE MONITORING Graphic representation of the pressure limit

- ALARMS / PRE-ALARMS / MANUAL PUSHER Protection for the ongoing infusion

- Green for infusion in progress, orange for pre-alarm and warnings, red for alarm.

- Visibility up to 5 meters. All alarms are expressed by means of light indicators, written words, pictograms and sound beep.

- INFUSION CONTROL Occlusion pressure pre-alarm, occlusion pressure alarm, end of (from 1 to 30 minutes) before the end of infusion, volume limit pre-alarm, volume limit alarm, keypad locking, hard and soft flow rate limits. low battery pre-alarm, discharged battery alarm, battery capacity display in hours and minutes, unconfirmed programming, technical malfunction alarm (auto-test , rotation), drive system advance check, watchdog check, communication connection failure.

- SCREEN Graphic LCD monochrome, size 66 mm x 33 mm ( 128 x 64 pixels).

- Battery backup – Minimum 10 Hrs

17.      Accessories (1) Power Cord

(2) Pole clamp