TECHNICAL SPECIFICATIONS-AUTOMATED BIOCHEMISTRY ANALYSER

1- Fully automatic random access clinical chemistry analyzer.

2- Assay type- end point, kinetic, turbidimetric, ISE, sample blank and reagent blank. Open reagent system.
3- Throughput- 180 test per hour or more than that (excluding ISE).
4- Separate reagent and sample probes with liquid level detector.

5- 40 or more reagent positions with on board refrigeration.

6- Integrated ISE unit for measurement of sodium, potassium and chloride.
7- Cuvette system: 40 or more reusable cuvettes with on board cuvette washing and temperature maintenance. 
8- Direct absorbance in cuvette (mono or bichromatic), filters ranging from 340 nm to 700 nm.

9- Minimum sample volume of 2 microliter or less than that.

10- Minimum reagent volume of 200 microliter or less than that.

11- Facility for test profiles, user defined order of analysis and calculated tests.

12- Manual key board entry for patient data as well as in built Bar code scanner for identification of samples with facility for routine, stat and quality control samples as well as to accommodate primary and secondary tubes.
13-  Windows based system with patient specific search option, facility for storage of >20,000 patient records and printing facility with external laser printer. A laser printer is to be supplied along with the equipment. Bidirectional interface with host computer.
14-  Facility for automatic calibration and quality control analysis.

15- Automatic start up facility(user defined).

16-  Facility for bar code generation. Two Bar code printers and software to be supplied along with the equipment.
17-  Technical specifications must be supported by manufacturer’s brochure failing which offer will be rejected. 

18- Vendor must be able to demonstrate the equipment at Hyderabad within specified time limit, if asked for; failing which offer will be rejected.

19-  Vendor must have a local office at Hyderabad with service engineer, ( at the time of submitting quotation, Address of the local office must be mentioned in the technical specifications), for prompt after sales service.
20-  If equipment or vendor does not meet any of the above mentioned conditions, offer will be rejected.
21- A power backup facility(UPS) of required capacity with two hour power backup is to be supplied along with the equipment.
